## Phone 1_GSM & WCDMA
Phone Performance Comparison Data
|  | GSM |  | WCDMA
Phone model: | RBT71UW | CMU200 setup |  | CMU200 setup
Manufacturer | RIM | Traffic Mode | Full Rate Version 2 | DCH Type | Speech
PIN | 21022414 | Bit Stream | Speech Codec/ Handset Low | Voice Source | Speechcodec Low
IMEI | 351846037105666 |  |  | Codec | AMR 12.2 kbps
FCC ID | L6ARBT70UW | [6 empty] | ;= reduced by the exclusion (one site only reporting this)
|  | Measured Avg. E-Field (V/m) | ANSI C63.19-2007 Calculations | [5 empty] | ANSI C63.19 Draft 1.13 Calculations | [5 empty] | 2007 v.s. Draft 1.13 Delta
|  |  | Measured PMF | Peak E-Field |  | M3 Limit | Margin to M3 Limit | M-Rating (2007) | Measured Avg. E-Field | Measured MIF | RF Audio Interference Level | M3 Limit | Margin to M3 Limit | M-Rating (draft 1.13)
Units |  | V/m | linear | V/m | dBV/m | dBV/m | dB |  | dBV/m | dB | dBV/m | dBV/m | dB |  | dB | Avg. E-Field deviation from the mean (dB)
Mode & Band | Frequency (MHz) | a | b | c = a * b | d = 20 log c | e | f = d - e |  | h = 20 log a | g | i = g + h | j | k = i - j |  | l = k - f
GSM 850 | 824.2 | 58.5 | 2.829 | 165.4965 | 44.3758 | 48.5 | -4.1242 | M3 | 35.3431 | 3.5 | 38.8431 | 45 | -6.1569 | M4 | -2.0327 | -0.012207
| 824.2 | 60.3 | 2.87 | 173.061 | 44.764 | 48.5 | -3.736 | M3 | 35.6063 | 3.52 | 39.1263 | 45 | -5.8737 | M4 | -2.1376 | 0.25102
| 824.2 | 56.59 | 2.86 | 161.8474 | 44.1821 | 48.5 | -4.3179 | M3 | 35.0548 | 3.5 | 38.5548 | 45 | -6.4452 | M4 | -2.1273 | -0.30053
| 824.2 | 59 | 3 | 177 | 44.9595 | 48.5 | -3.5405 | M3 | 35.417 |  |  | 45 |  |  |  | 0.061716
| 836.8 | 49.8 | 2.829 | 140.8842 | 42.9772 | 48.5 | -5.5228 | M4 | 33.9446 | 3.5 | 37.4446 | 45 | -7.5554 | M4 | -2.0327 | -0.82195
| 836.8 | 58 | 2.87 | 166.46 | 44.4262 | 48.5 | -4.0738 | M3 | 35.2686 | 3.52 | 38.7886 | 45 | -6.2114 | M4 | -2.1376 | 0.50202
| 836.8 | 53.24 | 2.86 | 152.2664 | 43.6521 | 48.5 | -4.8479 | M3 | 34.5248 | 3.5 | 38.0248 | 45 | -6.9752 | M4 | -2.1273 | -0.24178
| 836.8 | 58.4 | 3 | 175.2 | 44.8707 | 48.5 | -3.6293 | M3 | 35.3283 |  |  | 45 |  |  |  | 0.56172
| 848.8 | 49.5 | 2.829 | 140.0355 | 42.9248 | 48.5 | -5.5752 | M4 | 33.8921 | 3.5 | 37.3921 | 45 | -7.6079 | M4 | -2.0327 | -0.74716
| 848.8 | 57.5 | 2.87 | 165.025 | 44.351 | 48.5 | -4.149 | M3 | 35.1934 | 3.52 | 38.7134 | 45 | -6.2866 | M4 | -2.1376 | 0.55409
| 848.8 | 54.8 | 2.86 | 156.728 | 43.9029 | 48.5 | -4.5971 | M3 | 34.7756 | 3.5 | 38.2756 | 45 | -6.7244 | M4 | -2.1273 | 0.13634
| 848.8 | 54.3 | 2.99 | 162.357 | 44.2094 | 48.5 | -4.2906 | M3 | 34.696 |  |  | 45 |  |  |  | 0.056729
GSM 1900 | 1850.2 | 33 | 2.79 | 92.07 | 39.2824 | 38.5 | 0.78236 | M2 | 30.3703 | 3.3 | 33.6703 | 35 | -1.3297 | M3 | -2.1121 | -0.0142
| 1850.2 | 30.7 | 2.79 | 85.653 | 38.6549 | 38.5 | 0.15485 | M2 | 29.7428 | 3.52 | 33.2628 | 35 | -1.7372 | M3 | -1.8921 | -0.64171
| 1850.2 | 38.38 | 2.84 | 108.9992 | 40.7485 | 38.5 | 2.2485 | M2 | 31.6821 | 3.5 | 35.1821 | 35 | 0.1821 | M2 | -2.0664 | 1.2976
| 1850.2 | 30.7 | 2.92 | 89.644 | 39.0504 | 38.5 | 0.55042 | M2 | 29.7428 |  |  | 35 |  |  |  | -0.64171
| 1880 | 28.1 | 2.79 | 78.399 | 37.8862 | 38.5 | -0.61379 | M3 | 28.9741 | 3.3 | 32.2741 | 35 | -2.7259 | M3 | -2.1121 | -0.080926
| 1880 | 25.7 | 2.79 | 71.703 | 37.1107 | 38.5 | -1.3893 | M3 | 28.1987 | 3.52 | 31.7187 | 35 | -3.2813 | M3 | -1.8921 | -0.85639
| 1880 | 33.19 | 2.84 | 94.2596 | 39.4865 | 38.5 | 0.98651 | M2 | 30.4201 | 3.5 | 33.9201 | 35 | -1.0799 | M3 | -2.0664 | 1.3651
| 1880 | 27 | 2.9 | 78.3 | 37.8752 | 38.5 | -0.62476 | M3 | 28.6273 |  |  | 35 |  |  |  | -0.42778
| 1909.8 | 27.5 | 2.79 | 76.725 | 37.6987 | 38.5 | -0.80126 | M3 | 28.7867 | 3.3 | 32.0867 | 35 | -2.9133 | M3 | -2.1121 | 0.42745
| 1909.8 | 24.4 | 2.79 | 68.076 | 36.6599 | 38.5 | -1.8401 | M3 | 27.7478 | 3.52 | 31.2678 | 35 | -3.7322 | M3 | -1.8921 | -0.6114
| 1909.8 | 29.29 | 2.84 | 83.1836 | 38.4008 | 38.5 | -0.099246 | M3 | 29.3344 | 3.5 | 32.8344 | 35 | -2.1656 | M3 | -2.0664 | 0.97519
| 1909.8 | 23.9 | 2.89 | 69.071 | 36.7859 | 38.5 | -1.7141 | M3 | 27.568 |  |  | 35 |  |  |  | -0.79124
WCDMA 850 | 826.4 | 59.9 | 0.961 | 57.5639 | 35.203 | 51 | -15.797 | M4 | 35.5485 | -17 | 18.5485 | 45 | -26.4515 | M4 | -10.6545 | 0.16871
| 826.4 | 60.5 | 0.94 | 56.87 | 35.0977 | 51 | -15.9023 | M4 | 35.6351 | -19.94 | 15.6951 | 45 | -29.3049 | M4 | -13.4026 | 0.25528
| 826.4 | 56.09 | 0.99 | 55.5291 | 34.8904 | 51 | -16.1096 | M4 | 34.9777 | -20.7 | 14.2777 | 45 | -30.7223 | M4 | -14.6127 | -0.40212
| 826.4 | 58.6 | 0.975 | 57.135 | 35.138 | 51 | -15.862 | M4 | 35.358 |  |  | 45 |  |  |  | -0.021874
| 836.4 | 47.3 | 0.961 | 45.4553 | 33.1517 | 51 | -17.8483 | M4 | 33.4972 | -17 | 16.4972 | 45 | -28.5028 | M4 | -10.6545 | -1.2105
| 836.4 | 62.2 | 0.94 | 58.468 | 35.3384 | 51 | -15.6616 | M4 | 35.8758 | -19.94 | 15.9358 | 45 | -29.0642 | M4 | -13.4026 | 1.1681
| 836.4 | 54.64 | 0.99 | 54.0936 | 34.6629 | 51 | -16.3371 | M4 | 34.7502 | -21.1 | 13.6502 | 45 | -31.3498 | M4 | -15.0127 | 0.042466
| 836.4 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -34.7077
| 846.6 | 55.2 | 0.961 | 53.0472 | 34.4932 | 51 | -16.5068 | M4 | 34.8388 | -17 | 17.8388 | 45 | -27.1612 | M4 | -10.6545 | -1.1733
| 846.6 | 73.5 | 0.94 | 69.09 | 36.7883 | 51 | -14.2117 | M4 | 37.3257 | -19.94 | 17.3857 | 45 | -27.6143 | M4 | -13.4026 | 1.3137
| 846.6 | 62.17 | 0.99 | 61.5483 | 35.7843 | 51 | -15.2157 | M4 | 35.8716 | -21 | 14.8716 | 45 | -30.1284 | M4 | -14.9127 | -0.14043
| 846.6 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -36.012
WCDMA 1900 | 1852.4 | 38 | 0.972 | 36.936 | 31.349 | 41 | -9.651 | M4 | 31.5957 | -15.1 | 16.4957 | 35 | -18.5043 | M4 | -8.8533 | -0.0025894
| 1852.4 | 34.2 | 0.91 | 31.122 | 29.8613 | 41 | -11.1387 | M4 | 30.6805 | -20.15 | 10.5305 | 35 | -24.4695 | M4 | -13.3308 | -0.91774
| 1852.4 | 42.26 | 1.01 | 42.6826 | 32.605 | 41 | -8.395 | M4 | 32.5186 | -19.9 | 12.6186 | 35 | -22.3814 | M4 | -13.9864 | 0.92033
| 1852.4 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -31.5983
| 1880 | 34.5 | 0.972 | 33.534 | 30.5097 | 41 | -10.4903 | M4 | 30.7564 | -15.1 | 15.6564 | 35 | -19.3436 | M4 | -8.8533 | -0.37144
| 1880 | 32.9 | 0.91 | 29.939 | 29.5247 | 41 | -11.4753 | M4 | 30.3439 | -20.15 | 10.1939 | 35 | -24.8061 | M4 | -13.3308 | -0.78391
| 1880 | 41.13 | 1.01 | 41.5413 | 32.3696 | 41 | -8.6304 | M4 | 32.2832 | -19.9 | 12.3832 | 35 | -22.6168 | M4 | -13.9864 | 1.1553
| 1880 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -31.1278
| 1907.6 | 34.8 | 0.972 | 33.8256 | 30.5849 | 41 | -10.4151 | M4 | 30.8316 | -15.1 | 15.7316 | 35 | -19.2684 | M4 | -8.8533 | 0.23051
| 1907.6 | 29.2 | 0.91 | 26.572 | 28.4885 | 41 | -12.5115 | M4 | 29.3077 | -20.15 | 9.1577 | 35 | -25.8423 | M4 | -13.3308 | -1.2934
| 1907.6 | 38.3 | 1.01 | 38.683 | 31.7504 | 41 | -9.2496 | M4 | 31.664 | -20.1 | 11.564 | 35 | -23.436 | M4 | -14.1864 | 1.0629
| 1907.6 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -30.6011

## Phone 2_GSM & CDMA
Phone Performance Comparison Data
|  | GSM |  | CDMA FR |  | CDMA 1/8th
Phone model: | RCF71CW | CMU200 setup |  | CMU200 setup |  | CMU200 setup
Manufacturer | RIM | Traffic Mode | Full Rate Version 2 | Service | Loopback | Service | Speech
PIN | 30B1FAC5 | Bit Stream | Speech Codec/ Handset Low | Service Option | SO2 | Service Option | SO3
IMEI | A000000D798788 |  |  | Radio Config | RC1 | Radio Config | RC1 |  |  | ;= questionable scan results (not included in deviation from mean results in final column)
FCC ID | L6ARCF70CW |  |  | Frame Rate | Full | Frame Rate | 1/8 th |  |  | ;= reduced by the exclusion (one site only reporting this)
[6 empty] | Voice Coder | 8k EVRC (Low)
|  | Measured Avg. E-Field (V/m) | ANSI C63.19-2007 Calculations | [5 empty] | ANSI C63.19 Draft 1.13 Calculations | [5 empty] | 2007 v.s. Draft 1.13 Delta
|  |  | Measured PMF | Peak E-Field |  | M3 Limit | Margin to M3 Limit | M-Rating (2007) | Measured Avg. E-Field | Measured MIF | RF Audio Interference Level | M3 Limit | Margin to M3 Limit | M-Rating (draft 1.13)
Units |  | V/m | linear | V/m | dBV/m | dBV/m | dB |  | dBV/m | dB | dBV/m | dBV/m | dB |  | dB | Avg. E-Field deviation from the mean (dB)
Mode & Band | Frequency (MHz) | a | b | c = a * b | d = 20 log c | e | f = d - e |  | h = 20 log a | g | i = g + h | j | k = i - j |  | l = k - f
GSM 850 | 824.2 | 59.1 | 2.829 | 167.1939 | 44.4644 | 48.5 | -4.0356 | M3 | 35.4317 | 3.5 | 38.9317 | 45 | -6.0683 | M4 | -2.0327 | -0.57817
| 824.2 | 72.1 | 2.93 | 211.253 | 46.4961 | 48.5 | -2.0039 | M3 | 37.1587 | 3.55 | 40.7087 | 45 | -4.2913 | M3 | -2.2874 | 1.1488
| 824.2 | 57.84 | 2.86 | 165.4224 | 44.3719 | 48.5 | -4.1281 | M3 | 35.2446 | 3.5 | 38.7446 | 45 | -6.2554 | M4 | -2.1273 | -0.76535
| 824.2 | 64.6 | 3.01 | 194.446 | 45.776 | 48.5 | -2.724 | M3 | 36.2047 |  |  | 45 |  |  |  | 0.19473
| 836.8 | 59.2 | 2.829 | 167.4768 | 44.4791 | 48.5 | -4.0209 | M3 | 35.4464 | 3.5 | 38.9464 | 45 | -6.0536 | M4 | -2.0327 | -1.3348
| 836.8 | 75.7 | 2.93 | 221.801 | 46.9193 | 48.5 | -1.5807 | M3 | 37.5819 | 3.55 | 41.1319 | 45 | -3.8681 | M3 | -2.2874 | 0.80064
| 836.8 | 69.33 | 2.86 | 198.2838 | 45.9457 | 48.5 | -2.5543 | M3 | 36.8184 | 3.5 | 40.3184 | 45 | -4.6816 | M3 | -2.1273 | 0.037143
| 836.8 | 73.1 | 3.02 | 220.762 | 46.8785 | 48.5 | -1.6215 | M3 | 37.2783 |  |  | 45 |  |  |  | 0.49707
| 848.8 | 61 | 2.829 | 172.569 | 44.7393 | 48.5 | -3.7607 | M3 | 35.7066 | 3.5 | 39.2066 | 45 | -5.7934 | M4 | -2.0327 | -1.006
| 848.8 | 74 | 2.93 | 216.82 | 46.722 | 48.5 | -1.778 | M3 | 37.3846 | 3.55 | 40.9346 | 45 | -4.0654 | M3 | -2.2874 | 0.67206
| 848.8 | 69.54 | 2.86 | 198.8844 | 45.972 | 48.5 | -2.528 | M3 | 36.8447 | 3.5 | 40.3447 | 45 | -4.6553 | M3 | -2.1273 | 0.13212
| 848.8 | 70.1 | 3.02 | 211.702 | 46.5145 | 48.5 | -1.9855 | M3 | 36.9144 |  |  | 45 |  |  |  | 0.20179
GSM 1900 | 1850.2 | 22.6 | 2.79 | 63.054 | 35.9943 | 38.5 | -2.5057 | M3 | 27.0822 | 3.3 | 30.3822 | 35 | -4.6178 | M3 | -2.1121 | -0.11037
| 1850.2 | 21.2 | 2.92 | 61.904 | 35.8344 | 38.5 | -2.6656 | M3 | 26.5267 | 3.54 | 30.0667 | 35 | -4.9333 | M3 | -2.2677 | -0.66582
| 1850.2 | 24.8 | 2.84 | 70.432 | 36.9554 | 38.5 | -1.5446 | M3 | 27.889 | 3.5 | 31.389 | 35 | -3.611 | M3 | -2.0664 | 0.69649
| 1850.2 | 23.1 | 2.88 | 66.528 | 36.4601 | 38.5 | -2.0399 | M3 | 27.2722 |  |  | 35 |  |  |  | 0.0797
| 1880 | 21.5 | 2.79 | 59.985 | 35.5609 | 38.5 | -2.9391 | M3 | 26.6488 | 3.3 | 29.9488 | 35 | -5.0512 | M4 | -2.1121 | -0.34972
| 1880 | 21.6 | 2.92 | 63.072 | 35.9967 | 38.5 | -2.5033 | M3 | 26.6891 | 3.54 | 30.2291 | 35 | -4.7709 | M3 | -2.2677 | -0.30941
| 1880 | 24.13 | 2.84 | 68.5292 | 36.7175 | 38.5 | -1.7825 | M3 | 27.6511 | 3.5 | 31.1511 | 35 | -3.8489 | M3 | -2.0664 | 0.65266
| 1880 | 22.4 | 2.88 | 64.512 | 36.1928 | 38.5 | -2.3072 | M3 | 27.005 |  |  | 35 |  |  |  | 0.0064726
| 1909.8 | 19.9 | 2.79 | 55.521 | 34.8891 | 38.5 | -3.6109 | M3 | 25.9771 | 3.3 | 29.2771 | 35 | -5.7229 | M4 | -2.1121 | -0.051567
| 1909.8 | 18.7 | 2.92 | 54.604 | 34.7445 | 38.5 | -3.7555 | M3 | 25.4368 | 3.54 | 28.9768 | 35 | -6.0232 | M4 | -2.2677 | -0.5918
| 1909.8 | 20.55 | 2.84 | 58.362 | 35.3226 | 38.5 | -3.1774 | M3 | 26.2562 | 3.5 | 29.7562 | 35 | -5.2438 | M4 | -2.0664 | 0.22761
| 1909.8 | 21 | 2.87 | 60.27 | 35.602 | 38.5 | -2.898 | M3 | 26.4444 |  |  | 35 |  |  |  | 0.41576
Cellular CDMA Full Rate | 824.7 | 64.8 | 0.93 | 60.264 | 35.6012 | 51 | -15.3988 | M4 | 36.2315 | -19.25 | 16.9815 | 45 | -28.0185 | M4 | -12.6197 | -0.5667
| 824.7 | 77.4 | 0.96 | 74.304 | 37.4202 | 51 | -13.5798 | M4 | 37.7748 | -19.13 | 18.6448 | 45 | -26.3552 | M4 | -12.7754 | 0.97662
| 824.7 | 64.46 | 0.97 | 62.5262 | 35.9212 | 51 | -15.0788 | M4 | 36.1858 | -19.3 | 16.8858 | 45 | -28.1142 | M4 | -13.0354 | -0.61239
| 824.7 | 70.8 | 0.97 | 68.676 | 36.7361 | 51 | -14.2639 | M4 | 37.0007 |  |  | 45 |  |  |  | 0.20247
| 836.52 | 63.4 | 0.93 | 58.962 | 35.4114 | 51 | -15.5886 | M4 | 36.0418 | -19.25 | 16.7918 | 45 | -28.2082 | M4 | -12.6197 | -0.67007
| 836.52 | 74.8 | 0.96 | 71.808 | 37.1235 | 51 | -13.8765 | M4 | 37.478 | -19.13 | 18.348 | 45 | -26.652 | M4 | -12.7754 | 0.76617
| 836.52 | 66.55 | 0.97 | 64.5535 | 36.1984 | 51 | -14.8016 | M4 | 36.463 | -19.3 | 17.163 | 45 | -27.837 | M4 | -13.0354 | -0.2489
| 836.52 | 69.7 | 0.97 | 67.609 | 36.6001 | 51 | -14.3999 | M4 | 36.8647 |  |  | 45 |  |  |  | 0.1528
| 848.31 | 69.5 | 0.93 | 64.635 | 36.2094 | 51 | -14.7906 | M4 | 36.8397 | -19.25 | 17.5897 | 45 | -27.4103 | M4 | -12.6197 | -0.80397
| 848.52 | 83.1 | 0.96 | 79.776 | 38.0374 | 51 | -12.9626 | M4 | 38.392 | -19.13 | 19.262 | 45 | -25.738 | M4 | -12.7754 | 0.74835
| 848.31 | 75.58 | 0.97 | 73.3126 | 37.3036 | 51 | -13.6964 | M4 | 37.5681 | -19.3 | 18.2681 | 45 | -26.7319 | M4 | -13.0354 | -0.075531
| 848.31 | 77.4 | 0.97 | 75.078 | 37.5103 | 51 | -13.4897 | M4 | 37.7748 |  |  | 45 |  |  |  | 0.13115
Cellular CDMA 1/8 Rate | 824.7 | 27.9 | 2.67 | 74.493 | 37.4423 | 51 | -13.5577 | M4 | 28.9121 | 3.23 | 32.1421 | 45 | -12.8579 | M4 | 0.69977 | 0.23574
| 824.7 | 30.4 | 2.68 | 81.472 | 38.2202 | 51 | -12.7798 | M4 | 29.6575 | 3.21 | 32.8675 | 45 | -12.1325 | M4 | 0.6473 | 0.98113
| 824.7 | 24.65 | 2.79 | 68.7735 | 36.7484 | 51 | -14.2516 | M4 | 27.8363 | 3.2 | 31.0363 | 45 | -13.9637 | M4 | 0.28792 | -0.84
| 824.7 | 26 | 2.63 | 68.38 | 36.6986 | 51 | -14.3014 | M4 | 28.2995 |  |  | 45 |  |  |  | -0.37687
| 836.52 | 24.7 | 2.67 | 65.949 | 36.3842 | 51 | -14.6158 | M4 | 27.8539 | 3.23 | 31.0839 | 45 | -13.9161 | M4 | 0.69977 | -0.15645
| 836.52 | 27.1 | 2.68 | 72.628 | 37.2221 | 51 | -13.7779 | M4 | 28.6594 | 3.21 | 31.8694 | 45 | -13.1306 | M4 | 0.6473 | 0.649
| 836.52 | 24 | 2.79 | 66.96 | 36.5163 | 51 | -14.4837 | M4 | 27.6042 | 3.2 | 30.8042 | 45 | -14.1958 | M4 | 0.28792 | -0.40616
| 836.52 | 24.9 | 2.63 | 65.487 | 36.3231 | 51 | -14.6769 | M4 | 27.924 |  |  | 45 |  |  |  | -0.086397
| 848.31 | 27.6 | 2.67 | 73.692 | 37.3484 | 51 | -13.6516 | M4 | 28.8182 | 3.23 | 32.0482 | 45 | -12.9518 | M4 | 0.69977 | -0.80216
| 848.52 | 35.2 | 2.68 | 94.336 | 39.4935 | 51 | -11.5065 | M4 | 30.9309 | 3.21 | 34.1409 | 45 | -10.8591 | M4 | 0.6473 | 1.3105
| 848.31 | 28.15 | 2.79 | 78.5385 | 37.9017 | 51 | -13.0983 | M4 | 28.9896 | 3.2 | 32.1896 | 45 | -12.8104 | M4 | 0.28792 | -0.63077
| 848.31 | 30.7 | 2.65 | 81.355 | 38.2077 | 51 | -12.7923 | M4 | 29.7428 |  |  | 45 |  |  |  | 0.12242
PCS CDMA Full Rate | 1851.25 | 11.3 | 0.99 | 11.187 | 20.9743 | 41 | -20.0257 | M4 | 21.0616 | -19.16 | 1.9016 | 35 | -33.0984 | M4 | -13.0727 | -8.2401
| 1851.25 | 27.6 | 1 | 27.6 | 28.8182 | 41 | -12.1818 | M4 | 28.8182 | -19.3 | 9.5182 | 35 | -25.4818 | M4 | -13.3 | -0.48351
| 1851.25 | 29.71 | 1.03 | 30.6013 | 29.7148 | 41 | -11.2852 | M4 | 29.4581 | -19.2 | 10.2581 | 35 | -24.7419 | M4 | -13.4567 | 0.15636
| 1851.25 | 30.3 | 0.96 | 29.088 | 29.2743 | 41 | -11.7257 | M4 | 29.6289 |  |  | 35 |  |  |  | 0.32716
| 1880 | 26.2 | 0.99 | 25.938 | 28.2787 | 41 | -12.7213 | M4 | 28.366 | -19.16 | 9.206 | 35 | -25.794 | M4 | -13.0727 | -0.0073355
| 1880 | 25.5 | 1 | 25.5 | 28.1308 | 41 | -12.8692 | M4 | 28.1308 | -19.3 | 8.8308 | 35 | -26.1692 | M4 | -13.3 | -0.24256
| 1880 | 26.21 | 1.03 | 26.9963 | 28.6261 | 41 | -12.3739 | M4 | 28.3693 | -19.7 | 8.6693 | 35 | -26.3307 | M4 | -13.9567 | -0.0040209
| 1880 | 27 | 0.96 | 25.92 | 28.2727 | 41 | -12.7273 | M4 | 28.6273 |  |  | 35 |  |  |  | 0.25391
| 1908.75 | 26.6 | 0.99 | 26.334 | 28.4103 | 41 | -12.5897 | M4 | 28.4976 | -19.16 | 9.3376 | 35 | -25.6624 | M4 | -13.0727 | 0.34766
| 1908.75 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -28.15
| 1908.75 | 25 | 1.03 | 25.75 | 28.2155 | 41 | -12.7845 | M4 | 27.9588 | -19.4 | 8.5588 | 35 | -26.4412 | M4 | -13.6567 | -0.19117
| 1908.75 | 25.1 | 0.96 | 24.096 | 27.6389 | 41 | -13.3611 | M4 | 27.9935 |  |  | 35 |  |  |  | -0.15649
PCS CDMA 1/8 Rate | 1851.25 | 11 | 2.77 | 30.47 | 29.6774 | 41 | -11.3226 | M4 | 20.8279 | 3.26 | 24.0879 | 35 | -10.9121 | M4 | 0.4104 | -0.62566
| 1851.25 | 12.7 | 2.84 | 36.068 | 31.1424 | 41 | -9.8576 | M4 | 22.0761 | 3.24 | 25.3161 | 35 | -9.6839 | M4 | 0.17363 | 0.62256
| 1851.25 | 11.65 | 2.77 | 32.2705 | 30.1761 | 41 | -10.8239 | M4 | 21.3265 | 3.2 | 24.5265 | 35 | -10.4735 | M4 | 0.3504 | -0.127
| 1851.25 | 12 | 2.59 | 31.08 | 29.8496 | 41 | -11.1504 | M4 | 21.5836 |  |  | 35 |  |  |  | 0.13011
| 1880 | 12.4 | 2.77 | 34.348 | 30.718 | 41 | -10.282 | M4 | 21.8684 | 3.26 | 25.1284 | 35 | -9.8716 | M4 | 0.4104 | 1.0531
| 1880 | 11.6 | 2.84 | 32.944 | 30.3555 | 41 | -10.6445 | M4 | 21.2892 | 3.24 | 24.5292 | 35 | -10.4708 | M4 | 0.17363 | 0.47382
| 1880 | 10.02 | 2.77 | 27.7554 | 28.8669 | 41 | -12.1331 | M4 | 20.0174 | 3.2 | 23.2174 | 35 | -11.7826 | M4 | 0.3504 | -0.79799
| 1880 | 10.1 | 2.59 | 26.159 | 28.3524 | 41 | -12.6476 | M4 | 20.0864 |  |  | 35 |  |  |  | -0.72892
| 1908.75 | 9.66 | 2.77 | 26.7582 | 28.5491 | 41 | -12.4509 | M4 | 19.6995 | 3.26 | 22.9595 | 35 | -12.0405 | M4 | 0.4104 | -0.30203
| 1908.75 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -20.0016
| 1908.75 | 9.68 | 2.77 | 26.8136 | 28.5671 | 41 | -12.4329 | M4 | 19.7175 | 3.2 | 22.9175 | 35 | -12.0825 | M4 | 0.3504 | -0.28407
| 1908.75 | 10.7 | 2.59 | 27.713 | 28.8537 | 41 | -12.1463 | M4 | 20.5877 |  |  | 35 |  |  |  | 0.5861

## Phone 3_iDEN
Phone Performance Comparison Data
IDEN
Phone model: | RCD21IN
Manufacturer | RIM | Tx settings
PIN | 40245A36 | Mode | pseudo training mode
IMEI | 10000003636820 | Frame | 1/3 rd
FCC ID | L6ARCD20IN
|  | Measured Avg. E-Field (V/m) | ANSI C63.19-2007 Calculations | [5 empty] | ANSI C63.19 Draft 1.13 Calculations | [5 empty] | 2007 v.s. Draft 1.13 Delta
|  |  | Measured PMF | Peak E-Field |  | M3 Limit | Margin to M3 Limit | M-Rating (2007) | Measured Avg. E-Field | Measured MIF | RF Audio Interference Level | M3 Limit | Margin to M3 Limit | M-Rating (draft 1.13)
Units |  | V/m | linear | V/m | dBV/m | dBV/m | dB |  | dBV/m | dB | dBV/m | dBV/m | dB |  | dB | Avg. E-Field deviation from the mean (dB)
Mode & Band | Frequency (MHz) | a | b | c = a * b | d = 20 log c | e | f = d - e |  | h = 20 log a | g | i = g + h | j | k = i - j |  | l = k - f
iDEN 800 | 806.01 | 67.7 | 3.33 | 225.441 | 47.0607 | 51 | -3.9393 | M3 | 36.6118 | 1.51 | 38.1218 | 45 | -6.8782 | M4 | -2.9389 | 0.38035
| 813.5 | 62.7 | 3.33 | 208.791 | 46.3942 | 51 | -4.6058 | M3 | 35.9454 | 1.51 | 37.4554 | 45 | -7.5446 | M4 | -2.9389 | -0.28608
| 824.99 | 64.1 | 3.33 | 213.453 | 46.586 | 51 | -4.414 | M3 | 36.1372 | 1.51 | 37.6472 | 45 | -7.3528 | M4 | -2.9389 | -0.094268
iDEN 900 | 896 | 84 | 3.02 | 253.68 | 48.0857 | 51 | -2.9143 | M3 | 38.4856 | 1.44 | 39.9256 | 45 | -5.0744 | M4 | -2.1601 | 1.4274
| 898.5 | 69.4 | 3.02 | 209.588 | 46.4273 | 51 | -4.5727 | M3 | 36.8272 | 1.44 | 38.2672 | 45 | -6.7328 | M4 | -2.1601 | -0.23101
| 901 | 62.1 | 3.02 | 187.542 | 45.462 | 51 | -5.538 | M4 | 35.8618 | 1.44 | 37.3018 | 45 | -7.6982 | M4 | -2.1601 | -1.1964

## Phone 4_GSM & WCDMA
Phone Performance Comparison Data
|  |  | GSM |  | WCDMA
Phone model: | SGH-A797 |  | CMU200 setup |  | CMU200 setup |  |  | The red flip phone
Manufacturer | Samsung |  | Traffic Mode | Full Rate Version 2 | DCH Type | Speech
PIN | R8YS816258N |  | Bit Stream | Speech Codec/ Handset Low | Voice Source | Speechcodec Low
IMEI | 354621030030071 |  |  |  | Codec | AMR 12.2 kbps
FCC ID | A3LSGHA797 | [7 empty] | ;= reduced by the exclusion (one site only reporting this)
Slider Closed
|  | Slider | Measured Avg. E-Field (V/m) | ANSI C63.19-2007 Calculations | [5 empty] | ANSI C63.19 Draft 1.13 Calculations | [5 empty] | 2007 v.s. Draft 1.13 Delta
|  |  |  | Measured PMF | Peak E-Field |  | M3 Limit | Margin to M3 Limit | M-Rating (2007) | Measured Avg. E-Field | Measured MIF | RF Audio Interference Level | M3 Limit | Margin to M3 Limit | M-Rating (draft 1.13)
Units |  |  | V/m | linear | V/m | dBV/m | dBV/m | dB |  | dBV/m | dB | dBV/m | dBV/m | dB |  | dB | Avg. E-Field deviation from the mean (dB)
Mode & Band | Frequency (MHz) |  | a | b | c = a * b | d = 20 log c | e | f = d - e |  | h = 20 log a | g | i = g + h | j | k = i - j |  | l = k - f
GSM 850 | 824.2 | Closed | 36.1 | 2.829 | 102.1269 | 40.1828 | 48.5 | -8.3172 | M4 | 31.1501 | 3.47 | 34.6201 | 45 | -10.3799 | M4 | -2.0627 | -0.45346
| 824.2 | Closed | 39.6 | 2.83 | 112.068 | 40.9896 | 48.5 | -7.5104 | M4 | 31.9539 | 3.49 | 35.4439 | 45 | -9.5561 | M4 | -2.0457 | 0.3503
| 824.2 | Closed | 37.73 | 2.86 | 107.9078 | 40.6611 | 48.5 | -7.8389 | M4 | 31.5337 | 3.5 | 35.0337 | 45 | -9.9663 | M4 | -2.1273 | -0.069868
| 824.2 | Closed | 38.8 | 2.95 | 114.46 | 41.1731 | 48.5 | -7.3269 | M4 | 31.7766 |  |  | 45 |  |  |  | 0.17303
| 836.8 | Closed | 38 | 2.829 | 107.502 | 40.6283 | 48.5 | -7.8717 | M4 | 31.5957 | 3.47 | 35.0657 | 45 | -9.9343 | M4 | -2.0627 | -1.2746
| 836.8 | Closed | 48.5 | 2.83 | 137.255 | 42.7506 | 48.5 | -5.7494 | M4 | 33.7148 | 3.49 | 37.2048 | 45 | -7.7952 | M4 | -2.0457 | 0.84461
| 836.8 | Closed | 42.48 | 2.86 | 121.4928 | 41.691 | 48.5 | -6.809 | M4 | 32.5637 | 3.5 | 36.0637 | 45 | -8.9363 | M4 | -2.1273 | -0.30654
| 836.8 | Closed | 47.9 | 2.97 | 142.263 | 43.0618 | 48.5 | -5.4382 | M4 | 33.6067 |  |  | 45 |  |  |  | 0.73648
| 848.8 | Closed | 39.4 | 2.829 | 111.4626 | 40.9426 | 48.5 | -7.5574 | M4 | 31.9099 | 3.47 | 35.3799 | 45 | -9.6201 | M4 | -2.0627 | -1.334
| 848.8 | Closed | 51.8 | 2.83 | 146.594 | 43.3223 | 48.5 | -5.1777 | M4 | 34.2866 | 3.49 | 37.7766 | 45 | -7.2234 | M4 | -2.0457 | 1.0426
| 848.8 | Closed | 45.66 | 2.86 | 130.5876 | 42.318 | 48.5 | -6.182 | M4 | 33.1907 | 3.5 | 36.6907 | 45 | -8.3093 | M4 | -2.1273 | -0.053231
| 848.8 | Closed | 47.8 | 2.97 | 141.966 | 43.0437 | 48.5 | -5.4563 | M4 | 33.5886 |  |  | 45 |  |  |  | 0.34461
GSM 1900 | 1850.2 | Closed | 19.9 | 2.79 | 55.521 | 34.8891 | 38.5 | -3.6109 | M3 | 25.9771 | 3.27 | 29.2471 | 35 | -5.7529 | M4 | -2.1421 | -0.53243
| 1850.2 | Closed | 19.4 | 2.82 | 54.708 | 34.761 | 38.5 | -3.739 | M3 | 25.756 | 3.38 | 29.136 | 35 | -5.864 | M4 | -2.125 | -0.75345
| 1850.2 | Closed | 26.62 | 2.84 | 75.6008 | 37.5705 | 38.5 | -0.92947 | M3 | 28.5042 | 3.5 | 32.0042 | 35 | -2.9958 | M3 | -2.0664 | 1.9947
| 1850.2 | Closed | 19.5 | 2.87 | 55.965 | 34.9583 | 38.5 | -3.5417 | M3 | 25.8007 |  |  | 35 |  |  |  | -0.7088
| 1880 | Closed | 19.7 | 2.79 | 54.963 | 34.8014 | 38.5 | -3.6986 | M3 | 25.8893 | 3.27 | 29.1593 | 35 | -5.8407 | M4 | -2.1421 | -1.2013
| 1880 | Closed | 20.1 | 2.82 | 56.682 | 35.0689 | 38.5 | -3.4311 | M3 | 26.0639 | 3.38 | 29.4439 | 35 | -5.5561 | M4 | -2.125 | -1.0268
| 1880 | Closed | 30.34 | 2.84 | 86.1656 | 38.7067 | 38.5 | 0.20668 | M2 | 29.6403 | 3.5 | 33.1403 | 35 | -1.8597 | M3 | -2.0664 | 2.5496
| 1880 | Closed | 21.8 | 2.88 | 62.784 | 35.957 | 38.5 | -2.543 | M3 | 26.7691 |  |  | 35 |  |  |  | -0.32154
| 1909.8 | Closed | 25.1 | 2.79 | 70.029 | 36.9056 | 38.5 | -1.5944 | M3 | 27.9935 | 3.27 | 31.2635 | 35 | -3.7365 | M3 | -2.1421 | 0.10979
| 1909.8 | Closed | 21.5 | 2.82 | 60.63 | 35.6538 | 38.5 | -2.8462 | M3 | 26.6488 | 3.38 | 30.0288 | 35 | -4.9712 | M3 | -2.125 | -1.2349
| 1909.8 | Closed | 30.27 | 2.84 | 85.9668 | 38.6866 | 38.5 | 0.18662 | M2 | 29.6202 | 3.5 | 33.1202 | 35 | -1.8798 | M3 | -2.0664 | 1.7366
| 1909.8 | Closed | 23.1 | 2.88 | 66.528 | 36.4601 | 38.5 | -2.0399 | M3 | 27.2722 |  |  | 35 |  |  |  | -0.61144
WCDMA 850 | 826.4 | Closed | 38 | 0.961 | 36.518 | 31.2501 | 51 | -19.7499 | M4 | 31.5957 | -13.8 | 17.7957 | 45 | -27.2043 | M4 | -7.4545 | -0.1954
| 826.4 | Closed | 40.6 | 1.02 | 41.412 | 32.3425 | 51 | -18.6575 | M4 | 32.1705 | -20.01 | 12.1605 | 45 | -32.8395 | M4 | -14.182 | 0.37945
| 826.4 | Closed | 38.51 | 0.99 | 38.1249 | 31.6242 | 51 | -19.3758 | M4 | 31.7115 | -20.3 | 11.4115 | 45 | -33.5885 | M4 | -14.2127 | -0.079601
| 826.4 | Closed | 38.4 | 0.98 | 37.632 | 31.5111 | 51 | -19.4889 | M4 | 31.6866 |  |  | 45 |  |  |  | -0.10445
| 836.4 | Closed | 32.3 | 0.961 | 31.0403 | 29.8385 | 51 | -21.1615 | M4 | 30.1841 | -13.8 | 16.3841 | 45 | -28.6159 | M4 | -7.4545 | -0.90436
| 836.4 | Closed | 41.8 | 1.02 | 42.636 | 32.5955 | 51 | -18.4045 | M4 | 32.4235 | -20.01 | 12.4135 | 45 | -32.5865 | M4 | -14.182 | 1.3351
| 836.4 | Closed | 31.84 | 0.99 | 31.5216 | 29.9722 | 51 | -21.0278 | M4 | 30.0595 | -20.2 | 9.8595 | 45 | -35.1405 | M4 | -14.1127 | -1.029
| 836.4 | Closed | 38.4 | 0.98 | 37.632 | 31.5111 | 51 | -19.4889 | M4 | 31.6866 |  |  | 45 |  |  |  | 0.59821
| 846.6 | Closed | 42.3 | 0.961 | 40.6503 | 32.1813 | 51 | -18.8187 | M4 | 32.5268 | -13.8 | 18.7268 | 45 | -26.2732 | M4 | -7.4545 | -1.1133
| 846.6 | Closed | 57 | 1.02 | 58.14 | 35.2895 | 51 | -15.7105 | M4 | 35.1175 | -20.01 | 15.1075 | 45 | -29.8925 | M4 | -14.182 | 1.4774
| 846.6 | Closed | 46 | 0.99 | 45.54 | 33.1679 | 51 | -17.8321 | M4 | 33.2552 | -21.5 | 11.7552 | 45 | -33.2448 | M4 | -15.4127 | -0.38494
| 846.6 | Closed | 48.2 | 0.98 | 47.236 | 33.4855 | 51 | -17.5145 | M4 | 33.6609 |  |  | 45 |  |  |  | 0.02084
WCDMA 1900 | 1852.4 | Closed | 30.9 | 0.972 | 30.0348 | 29.5525 | 41 | -11.4475 | M4 | 29.7992 | -13.9 | 15.8992 | 35 | -19.1008 | M4 | -7.6533 | -0.36578
| 1852.4 | Closed | 31.2 | 1.03 | 32.136 | 30.1398 | 41 | -10.8602 | M4 | 29.8831 | -20.13 | 9.7531 | 35 | -25.2469 | M4 | -14.3867 | -0.28185
| 1852.4 | Closed | 39.13 | 1.01 | 39.5213 | 31.9366 | 41 | -9.0634 | M4 | 31.8502 | -20.3 | 11.5502 | 35 | -23.4498 | M4 | -14.3864 | 1.6853
| 1852.4 | Closed | 28.6 | 0.99 | 28.314 | 29.04 | 41 | -11.96 | M4 | 29.1273 |  |  | 35 |  |  |  | -1.0376
| 1880 | Closed | 27.6 | 0.972 | 26.8272 | 28.5715 | 41 | -12.4285 | M4 | 28.8182 | -13.9 | 14.9182 | 35 | -20.0818 | M4 | -7.6533 | -0.42679
| 1880 | Closed | 27.2 | 1.03 | 28.016 | 28.9481 | 41 | -12.0519 | M4 | 28.6914 | -20.13 | 8.5614 | 35 | -26.4386 | M4 | -14.3867 | -0.5536
| 1880 | Closed | 35.37 | 1.01 | 35.7237 | 31.0591 | 41 | -9.9409 | M4 | 30.9727 | -20.8 | 10.1727 | 35 | -24.8273 | M4 | -14.8864 | 1.7277
| 1880 | Closed | 26.6 | 0.99 | 26.334 | 28.4103 | 41 | -12.5897 | M4 | 28.4976 |  |  | 35 |  |  |  | -0.74734
| 1907.6 | Closed | 29.1 | 0.972 | 28.2852 | 29.0312 | 41 | -11.9688 | M4 | 29.2779 | -13.9 | 15.3779 | 35 | -19.6221 | M4 | -7.6533 | 0.097219
| 1907.6 | Closed | 26.5 | 1.03 | 27.295 | 28.7217 | 41 | -12.2783 | M4 | 28.4649 | -20.13 | 8.3349 | 35 | -26.6651 | M4 | -14.3867 | -0.71572
| 1907.6 | Closed | 35.71 | 1.01 | 36.0671 | 31.1422 | 41 | -9.8578 | M4 | 31.0558 | -21.8 | 9.2558 | 35 | -25.7442 | M4 | -15.8864 | 1.8752
| 1907.6 | Closed | 24.9 | 1 | 24.9 | 27.924 | 41 | -13.076 | M4 | 27.924 |  |  | 35 |  |  |  | -1.2567
GSM 850 | 824.2 | Open | 43.7 | 2.829 | 123.6273 | 41.8423 | 48.5 | -6.6577 | M4 | 32.8096 | 3.47 | 36.2796 | 45 | -8.7204 | M4 | -2.0627 | -0.39769
| 824.2 | Open | 51.9 | 2.83 | 146.877 | 43.3391 | 48.5 | -5.1609 | M4 | 34.3033 | 3.49 | 37.7933 | 45 | -7.2067 | M4 | -2.0457 | 1.096
| 824.2 | Open | 43.01 | 2.86 | 123.0086 | 41.7987 | 48.5 | -6.7013 | M4 | 32.6714 | 3.5 | 36.1714 | 45 | -8.8286 | M4 | -2.1273 | -0.53593
| 824.2 | Open | 44.9 | 2.96 | 132.904 | 42.4708 | 48.5 | -6.0292 | M4 | 33.0449 |  |  | 45 |  |  |  | -0.1624
| 836.8 | Open | 47.9 | 2.829 | 135.5091 | 42.6394 | 48.5 | -5.8606 | M4 | 33.6067 | 3.47 | 37.0767 | 45 | -7.9233 | M4 | -2.0627 | -1.1699
| 836.8 | Open | 63.4 | 2.83 | 179.422 | 45.0775 | 48.5 | -3.4225 | M3 | 36.0418 | 3.49 | 39.5318 | 45 | -5.4682 | M4 | -2.0457 | 1.2652
| 836.8 | Open | 53.92 | 2.86 | 154.2112 | 43.7623 | 48.5 | -4.7377 | M3 | 34.635 | 3.5 | 38.135 | 45 | -6.865 | M4 | -2.1273 | -0.14163
| 836.8 | Open | 55.1 | 2.99 | 164.749 | 44.3365 | 48.5 | -4.1635 | M3 | 34.823 |  |  | 45 |  |  |  | 0.046401
| 848.8 | Open | 52.8 | 2.829 | 149.3712 | 43.4853 | 48.5 | -5.0147 | M4 | 34.4527 | 3.47 | 37.9227 | 45 | -7.0773 | M4 | -2.0627 | -0.9347
| 848.8 | Open | 67.2 | 2.83 | 190.176 | 45.5831 | 48.5 | -2.9169 | M3 | 36.5474 | 3.49 | 40.0374 | 45 | -4.9626 | M3 | -2.0457 | 1.16
| 848.8 | Open | 59.73 | 2.86 | 170.8278 | 44.6512 | 48.5 | -3.8488 | M3 | 35.5239 | 3.5 | 39.0239 | 45 | -5.9761 | M4 | -2.1273 | 0.13648
| 848.8 | Open | 56.4 | 2.99 | 168.636 | 44.539 | 48.5 | -3.961 | M3 | 35.0256 |  |  | 45 |  |  |  | -0.36179
GSM 1900 | 1850.2 | Open | 18.4 | 2.79 | 51.336 | 34.2084 | 38.5 | -4.2916 | M3 | 25.2964 | 3.27 | 28.5664 | 35 | -6.4336 | M4 | -2.1421 | -0.78598
| 1850.2 | Open | 17.9 | 2.82 | 50.478 | 34.062 | 38.5 | -4.438 | M3 | 25.0571 | 3.38 | 28.4371 | 35 | -6.5629 | M4 | -2.125 | -1.0253
| 1850.2 | Open | 25.5 | 2.84 | 72.42 | 37.1972 | 38.5 | -1.3028 | M3 | 28.1308 | 3.5 | 31.6308 | 35 | -3.3692 | M3 | -2.0664 | 2.0485
| 1850.2 | Open | 19.6 | 2.87 | 56.252 | 35.0028 | 38.5 | -3.4972 | M3 | 25.8451 |  |  | 35 |  |  |  | -0.23721
| 1880 | Open | 20.8 | 2.79 | 58.032 | 35.2734 | 38.5 | -3.2266 | M3 | 26.3613 | 3.27 | 29.6313 | 35 | -5.3687 | M4 | -2.1421 | -0.84425
| 1880 | Open | 20.7 | 2.82 | 58.374 | 35.3244 | 38.5 | -3.1756 | M3 | 26.3194 | 3.38 | 29.6994 | 35 | -5.3006 | M4 | -2.125 | -0.88611
| 1880 | Open | 28.63 | 2.84 | 81.3092 | 38.2028 | 38.5 | -0.29721 | M3 | 29.1364 | 3.5 | 32.6364 | 35 | -2.3636 | M3 | -2.0664 | 1.9309
| 1880 | Open | 22.4 | 2.88 | 64.512 | 36.1928 | 38.5 | -2.3072 | M3 | 27.005 |  |  | 35 |  |  |  | -0.20055
| 1909.8 | Open | 23.8 | 2.79 | 66.402 | 36.4436 | 38.5 | -2.0564 | M3 | 27.5315 | 3.27 | 30.8015 | 35 | -4.1985 | M3 | -2.1421 | -0.27189
| 1909.8 | Open | 22.2 | 2.82 | 62.604 | 35.932 | 38.5 | -2.568 | M3 | 26.9271 | 3.38 | 30.3071 | 35 | -4.6929 | M3 | -2.125 | -0.87637
| 1909.8 | Open | 30.32 | 2.84 | 86.1088 | 38.701 | 38.5 | 0.20095 | M2 | 29.6346 | 3.5 | 33.1346 | 35 | -1.8654 | M3 | -2.0664 | 1.8312
| 1909.8 | Open | 22.7 | 2.88 | 65.376 | 36.3084 | 38.5 | -2.1916 | M3 | 27.1205 |  |  | 35 |  |  |  | -0.68291
WCDMA 850 | 826.4 | Open | 52.5 | 0.961 | 50.4525 | 34.0577 | 51 | -16.9423 | M4 | 34.4032 | -13.8 | 20.6032 | 45 | -24.3968 | M4 | -7.4545 | 0.35261
| 826.4 | Open | 53.4 | 1.02 | 54.468 | 34.7228 | 51 | -16.2772 | M4 | 34.5508 | -20.01 | 14.5408 | 45 | -30.4592 | M4 | -14.182 | 0.50025
| 826.4 | Open | 47.4 | 0.99 | 46.926 | 33.4283 | 51 | -17.5717 | M4 | 33.5156 | -20.3 | 13.2156 | 45 | -31.7844 | M4 | -14.2127 | -0.53501
| 826.4 | Open | 48.6 | 0.98 | 47.628 | 33.5572 | 51 | -17.4428 | M4 | 33.7327 |  |  | 45 |  |  |  | -0.31785
| 836.4 | Open | 39.7 | 0.961 | 38.1517 | 31.6303 | 51 | -19.3697 | M4 | 31.9758 | -13.8 | 18.1758 | 45 | -26.8242 | M4 | -7.4545 | -1.3416
| 836.4 | Open | 55.6 | 1.02 | 56.712 | 35.0735 | 51 | -15.9265 | M4 | 34.9015 | -20.01 | 14.8915 | 45 | -30.1085 | M4 | -14.182 | 1.584
| 836.4 | Open | 44.7 | 0.99 | 44.253 | 32.9189 | 51 | -18.0811 | M4 | 33.0062 | -20.2 | 12.8062 | 45 | -32.1938 | M4 | -14.1127 | -0.3113
| 836.4 | Open | 46.7 | 0.98 | 45.766 | 33.2109 | 51 | -17.7891 | M4 | 33.3863 |  |  | 45 |  |  |  | 0.068889
| 846.6 | Open | 53.4 | 0.961 | 51.3174 | 34.2053 | 51 | -16.7947 | M4 | 34.5508 | -13.8 | 20.7508 | 45 | -24.2492 | M4 | -7.4545 | -0.8023
| 846.6 | Open | 70.5 | 1.02 | 71.91 | 37.1358 | 51 | -13.8642 | M4 | 36.9638 | -20.01 | 16.9538 | 45 | -28.0462 | M4 | -14.182 | 1.6107
| 846.6 | Open | 56.11 | 0.99 | 55.5489 | 34.8935 | 51 | -16.1065 | M4 | 34.9808 | -21.5 | 13.4808 | 45 | -31.5192 | M4 | -15.4127 | -0.37232
| 846.6 | Open | 55.7 | 0.98 | 54.586 | 34.7416 | 51 | -16.2584 | M4 | 34.9171 |  |  | 45 |  |  |  | -0.43603
WCDMA 1900 | 1852.4 | Open | 29.9 | 0.972 | 29.0628 | 29.2667 | 41 | -11.7333 | M4 | 29.5134 | -13.9 | 15.6134 | 35 | -19.3866 | M4 | -7.6533 | -0.29158
| 1852.4 | Open | 30 | 1.03 | 30.9 | 29.7992 | 41 | -11.2008 | M4 | 29.5424 | -20.13 | 9.4124 | 35 | -25.5876 | M4 | -14.3867 | -0.26258
| 1852.4 | Open | 34.9 | 1.01 | 35.249 | 30.9429 | 41 | -10.0571 | M4 | 30.8565 | -20.3 | 10.5565 | 35 | -24.4435 | M4 | -14.3864 | 1.0515
| 1852.4 | Open | 29.2 | 0.99 | 28.908 | 29.2204 | 41 | -11.7796 | M4 | 29.3077 |  |  | 35 |  |  |  | -0.49735
| 1880 | Open | 29.2 | 0.972 | 28.3824 | 29.061 | 41 | -11.939 | M4 | 29.3077 | -13.9 | 15.4077 | 35 | -19.5923 | M4 | -7.6533 | -0.096229
| 1880 | Open | 28.2 | 1.03 | 29.046 | 29.2617 | 41 | -11.7383 | M4 | 29.005 | -20.13 | 8.875 | 35 | -26.125 | M4 | -14.3867 | -0.3989
| 1880 | Open | 32.96 | 1.01 | 33.2896 | 30.4462 | 41 | -10.5538 | M4 | 30.3597 | -20.8 | 9.5597 | 35 | -25.4403 | M4 | -14.8864 | 0.95586
| 1880 | Open | 28 | 0.99 | 27.72 | 28.8559 | 41 | -12.1441 | M4 | 28.9432 |  |  | 35 |  |  |  | -0.46073
| 1907.6 | Open | 27.8 | 0.972 | 27.0216 | 28.6342 | 41 | -12.3658 | M4 | 28.8809 | -13.9 | 14.9809 | 35 | -20.0191 | M4 | -7.6533 | -0.080843
| 1907.6 | Open | 25.9 | 1.03 | 26.677 | 28.5227 | 41 | -12.4773 | M4 | 28.266 | -20.13 | 8.136 | 35 | -26.864 | M4 | -14.3867 | -0.69574
| 1907.6 | Open | 34.44 | 1.01 | 34.7844 | 30.8277 | 41 | -10.1723 | M4 | 30.7413 | -21.8 | 8.9413 | 35 | -26.0587 | M4 | -15.8864 | 1.7795
| 1907.6 | Open | 25 | 1 | 25 | 27.9588 | 41 | -13.0412 | M4 | 27.9588 |  |  | 35 |  |  |  | -1.0029

## Phone 5_CDMA
Phone Performance Comparison Data
|  | CDMA FR |  | CDMA 1/8th
Phone model: | SCH-U450 | CMU200 setup |  | CMU200 setup
Manufacturer | Samsung | Service | Loopback | Service | Speech
S/N | 8010D924 | Service Option | SO2 | Service Option | SO3 |  |  | SID : 331, NID 1
MEID | A00000177FC7CE | Radio Config | RC1 | Radio Config | RC1
FCC ID | A3LSCHU450 | Frame Rate | Full | Frame Rate | 1/8 th |  |  | ;= reduced by the exclusion (one site only reporting this)
|  |  |  | Voice Coder | 8k EVRC (Low)
|  | Measured Avg. E-Field (V/m) | ANSI C63.19-2007 Calculations | [5 empty] | ANSI C63.19 Draft 1.13 Calculations | [5 empty] | 2007 v.s. Draft 1.13 Delta
|  |  | Measured PMF | Peak E-Field |  | M3 Limit | Margin to M3 Limit | M-Rating (2007) | Measured Avg. E-Field | Measured MIF | RF Audio Interference Level | M3 Limit | Margin to M3 Limit | M-Rating (draft 1.13)
Units |  | V/m | linear | V/m | dBV/m | dBV/m | dB |  | dBV/m | dB | dBV/m | dBV/m | dB |  | dB | Avg. E-Field deviation from the mean (dB)
Mode & Band | Frequency (MHz) | a | b | c = a * b | d = 20 log c | e | f = d - e |  | h = 20 log a | g | i = g + h | j | k = i - j |  | l = k - f
Cellular CDMA Full Rate | 824.7 | 69 | 0.928 | 64.032 | 36.1279 | 51 | -14.8721 | M4 | 36.777 | -18.3 | 18.477 | 45 | -26.523 | M4 | -11.651 | 0.04492
| 824.7 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -36.7321
| 824.7 | 68.29 | 0.97 | 66.2413 | 36.4226 | 51 | -14.5774 | M4 | 36.6871 | -19.4 | 17.2871 | 45 | -27.7129 | M4 | -13.1354 | -0.04492
| 824.7 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -36.7321
| 836.52 | 59.4 | 0.928 | 55.1232 | 34.8267 | 51 | -16.1733 | M4 | 35.4757 | -18.3 | 17.1757 | 45 | -27.8243 | M4 | -11.651 | -0.2105
| 836.52 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -35.6862
| 836.52 | 62.35 | 0.97 | 60.4795 | 35.6322 | 51 | -15.3678 | M4 | 35.8967 | -19.1 | 16.7967 | 45 | -28.2033 | M4 | -12.8354 | 0.2105
| 836.52 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -35.6862
| 848.31 | 54.8 | 0.928 | 50.8544 | 34.1266 | 51 | -16.8734 | M4 | 34.7756 | -18.3 | 16.4756 | 45 | -28.5244 | M4 | -11.651 | -0.20887
| 848.31 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -34.9845
| 848.31 | 57.5 | 0.97 | 55.775 | 34.9288 | 51 | -16.0712 | M4 | 35.1934 | -19.2 | 15.9934 | 45 | -29.0066 | M4 | -12.9354 | 0.20887
| 848.31 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -34.9845
Cellular CDMA 1/8 Rate | 824.7 | 24.6 | 2.67 | 65.682 | 36.3489 | 51 | -14.6511 | M4 | 27.8187 | 3.27 | 31.0887 | 45 | -13.9113 | M4 | 0.73977 | 0.035453
| 824.7 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -27.7832
| 824.7 | 24.4 | 2.79 | 68.076 | 36.6599 | 51 | -14.3401 | M4 | 27.7478 | 3.2 | 30.9478 | 45 | -14.0522 | M4 | 0.28792 | -0.035453
| 824.7 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -27.7832
| 836.52 | 22.3 | 2.67 | 59.541 | 35.4963 | 51 | -15.5037 | M4 | 26.9661 | 3.27 | 30.2361 | 45 | -14.7639 | M4 | 0.73977 | -0.16809
| 836.52 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -27.1342
| 836.52 | 23.18 | 2.79 | 64.6722 | 36.2144 | 51 | -14.7856 | M4 | 27.3023 | 3.2 | 30.5023 | 45 | -14.4977 | M4 | 0.28792 | 0.16809
| 836.52 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -27.1342
| 848.31 | 20.9 | 2.67 | 55.803 | 34.9332 | 51 | -16.0668 | M4 | 26.4029 | 3.27 | 29.6729 | 45 | -15.3271 | M4 | 0.73977 | -0.039303
| 848.31 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -26.4422
| 848.31 | 21.09 | 2.79 | 58.8411 | 35.3936 | 51 | -15.6064 | M4 | 26.4815 | 3.2 | 29.6815 | 45 | -15.3185 | M4 | 0.28792 | 0.039303
| 848.31 |  |  |  |  | 51 | [5 empty] | 45 |  |  |  | -26.4422
PCS CDMA Full Rate | 1851.25 | 32.5 | 0.99 | 32.175 | 30.1504 | 41 | -10.8496 | M4 | 30.2377 | -19.36 | 10.8777 | 35 | -24.1223 | M4 | -13.2727 | -0.43453
| 1851.25 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -30.6722
| 1851.25 | 35.92 | 1.03 | 36.9976 | 31.3635 | 41 | -9.6365 | M4 | 31.1067 | -19.4 | 11.7067 | 35 | -23.2933 | M4 | -13.6567 | 0.43453
| 1851.25 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -30.6722
| 1880 | 25.1 | 0.99 | 24.849 | 27.9062 | 41 | -13.0938 | M4 | 27.9935 | -19.36 | 8.6335 | 35 | -26.3665 | M4 | -13.2727 | -0.43629
| 1880 | 23.6 | 1.01 | 23.836 | 27.5447 | 41 | -13.4553 | M4 | 27.4582 | -18.5 | 8.9582 | 35 | -26.0418 | M4 | -12.5864 | -0.97152
| 1880 | 30.68 | 1.03 | 31.6004 | 29.9939 | 41 | -11.0061 | M4 | 29.7371 | -19.6 | 10.1371 | 35 | -24.8629 | M4 | -13.8567 | 1.3073
| 1880 | 26.7 | 0.95 | 25.365 | 28.0847 | 41 | -12.9153 | M4 | 28.5302 |  |  | 35 |  |  |  | 0.10046
| 1908.75 | 23.8 | 0.99 | 23.562 | 27.4442 | 41 | -13.5558 | M4 | 27.5315 | -19.36 | 8.1715 | 35 | -26.8285 | M4 | -13.2727 | -0.85072
| 1908.75 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -28.3823
| 1908.75 | 28.95 | 1.03 | 29.8185 | 29.4897 | 41 | -11.5103 | M4 | 29.233 | -19.6 | 9.633 | 35 | -25.367 | M4 | -13.8567 | 0.85072
| 1908.75 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -28.3823
PCS CDMA 1/8 Rate | 1851.25 | 11.2 | 2.765 | 30.968 | 29.8183 | 41 | -11.1817 | M4 | 20.9844 | 3.24 | 24.2244 | 35 | -10.7756 | M4 | 0.4061 | -1.2192
| 1851.25 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -22.2036
| 1851.25 | 14.83 | 2.77 | 41.0791 | 32.2724 | 41 | -8.7276 | M4 | 23.4228 | 3.2 | 26.6228 | 35 | -8.3772 | M4 | 0.3504 | 1.2192
| 1851.25 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -22.2036
| 1880 | 8.9 | 2.765 | 24.6085 | 27.8217 | 41 | -13.1783 | M4 | 18.9878 | 3.24 | 22.2278 | 35 | -12.7722 | M4 | 0.4061 | -1.0101
| 1880 | 9.11 | 2.866 | 26.1093 | 28.3359 | 41 | -12.6641 | M4 | 19.1904 | 3.15 | 22.3404 | 35 | -12.6596 | M4 | 0.0044763 | -0.80751
| 1880 | 12.08 | 2.77 | 33.4616 | 30.4909 | 41 | -10.5091 | M4 | 21.6413 | 3.2 | 24.8413 | 35 | -10.1587 | M4 | 0.3504 | 1.6435
| 1880 | 10.2 | 2.57 | 26.214 | 28.3707 | 41 | -12.6293 | M4 | 20.172 |  |  | 35 |  |  |  | 0.17413
| 1908.75 | 8.18 | 2.765 | 22.6177 | 27.089 | 41 | -13.911 | M4 | 18.2551 | 3.24 | 21.4951 | 35 | -13.5049 | M4 | 0.4061 | -1.2188
| 1908.75 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -19.4738
| 1908.75 | 10.83 | 2.77 | 29.9991 | 29.5422 | 41 | -11.4578 | M4 | 20.6926 | 3.2 | 23.8926 | 35 | -11.1074 | M4 | 0.3504 | 1.2188
| 1908.75 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -19.4738

## Phone 6_CDMA
Phone Performance Comparison Data
|  | CDMA FR |  | CDMA 1/8th
Phone model: | SCH-U750 | CMU200 setup |  | CMU200 setup
Manufacturer | Samsung | Service | Loopback | Service | Speech
SKU | SCHU750HAV | Service Option | SO2 | Service Option | SO3
MEID | A0000017B0328D | Radio Config | RC1 | Radio Config | RC1
FCC ID | A3LSCHU750 | Frame Rate | Full | Frame Rate | 1/8 th
|  |  |  | Voice Coder | 8k EVRC (Low)
|  | Measured Avg. E-Field (V/m) | ANSI C63.19-2007 Calculations | [5 empty] | ANSI C63.19 Draft 1.13 Calculations | [5 empty] | 2007 v.s. Draft 1.13 Delta
|  |  | Measured PMF | Peak E-Field |  | M3 Limit | Margin to M3 Limit | M-Rating (2007) | Measured Avg. E-Field | Measured MIF | RF Audio Interference Level | M3 Limit | Margin to M3 Limit | M-Rating (draft 1.13)
Units |  | V/m | linear | V/m | dBV/m | dBV/m | dB |  | dBV/m | dB | dBV/m | dBV/m | dB |  | dB | Avg. E-Field deviation from the mean (dB)
Mode & Band | Frequency (MHz) | a | b | c = a * b | d = 20 log c | e | f = d - e |  | h = 20 log a | g | i = g + h | j | k = i - j |  | l = k - f
Cellular CDMA Full Rate | 824.7 | 26.7 | 0.928 | 24.7776 | 27.8812 | 51 | -23.1188 | M4 | 28.5302 | -19.12 | 9.4102 | 45 | -35.5898 | M4 | -12.471 | -0.66645
| 824.7 | 28.1 | 0.928 | 26.0768 | 28.3251 | 51 | -22.6749 | M4 | 28.9741 | -19.12 | 9.8541 | 45 | -35.1459 | M4 | -12.471 | -0.22255
| 824.7 | 33.48 | 0.97 | 32.4756 | 30.2311 | 51 | -20.7689 | M4 | 30.4957 | -19.1 | 11.3957 | 45 | -33.6043 | M4 | -12.8354 | 1.299
| 824.7 | 27.5 | 0.96 | 26.4 | 28.4321 | 51 | -22.5679 | M4 | 28.7867 |  |  | 45 |  |  |  | -0.41002
| 836.52 | 28.4 | 0.928 | 26.3552 | 28.4173 | 51 | -22.5827 | M4 | 29.0664 | -19.12 | 9.9464 | 45 | -35.0536 | M4 | -12.471 | -0.8547
| 836.52 | 33.7 | 0.928 | 31.2736 | 29.9036 | 51 | -21.0964 | M4 | 30.5526 | -19.12 | 11.4326 | 45 | -33.5674 | M4 | -12.471 | 0.63153
| 836.52 | 32.19 | 0.97 | 31.2243 | 29.8899 | 51 | -21.1101 | M4 | 30.1544 | -19.1 | 11.0544 | 45 | -33.9456 | M4 | -12.8354 | 0.23335
| 836.52 | 31.3 | 0.96 | 30.048 | 29.5563 | 51 | -21.4437 | M4 | 29.9109 |  |  | 45 |  |  |  | -0.010181
| 848.31 | 31 | 0.928 | 28.768 | 29.1782 | 51 | -21.8218 | M4 | 29.8272 | -19.12 | 10.7072 | 45 | -34.2928 | M4 | -12.471 | -1.0867
| 848.52 | 41.5 | 0.928 | 38.512 | 31.7119 | 51 | -19.2881 | M4 | 32.361 | -19.12 | 13.241 | 45 | -31.759 | M4 | -12.471 | 1.447
| 848.31 | 32.8 | 0.97 | 31.816 | 30.0529 | 51 | -20.9471 | M4 | 30.3175 | -19.2 | 11.1175 | 45 | -33.8825 | M4 | -12.9354 | -0.59648
| 848.31 | 36.1 | 0.96 | 34.656 | 30.7956 | 51 | -20.2044 | M4 | 31.1501 |  |  | 45 |  |  |  | 0.23619
Cellular CDMA 1/8 Rate | 824.7 | 9 | 2.67 | 24.03 | 27.6151 | 51 | -23.3849 | M4 | 19.0849 | 3.29 | 22.3749 | 45 | -22.6251 | M4 | 0.75977 | -1.1804
| 824.7 | 11 | 2.81 | 30.91 | 29.802 | 51 | -21.198 | M4 | 20.8279 | 3.28 | 24.1079 | 45 | -20.8921 | M4 | 0.30587 | 0.56258
| 824.7 | 10.87 | 2.79 | 30.3273 | 29.6367 | 51 | -21.3633 | M4 | 20.7246 | 3.2 | 23.9246 | 45 | -21.0754 | M4 | 0.28792 | 0.45932
| 824.7 | 10.5 | 2.57 | 26.985 | 28.6224 | 51 | -22.3776 | M4 | 20.4238 |  |  | 45 |  |  |  | 0.15852
| 836.52 | 10 | 2.67 | 26.7 | 28.5302 | 51 | -22.4698 | M4 | 20 | 3.29 | 23.29 | 45 | -21.71 | M4 | 0.75977 | -1.0864
| 836.52 | 13.7 | 2.81 | 38.497 | 31.7085 | 51 | -19.2915 | M4 | 22.7344 | 3.28 | 26.0144 | 45 | -18.9856 | M4 | 0.30587 | 1.648
| 836.52 | 10.56 | 2.79 | 29.4624 | 29.3854 | 51 | -21.6146 | M4 | 20.4733 | 3.2 | 23.6733 | 45 | -21.3267 | M4 | 0.28792 | -0.61317
| 836.52 | 11.4 | 2.57 | 29.298 | 29.3368 | 51 | -21.6632 | M4 | 21.1381 |  |  | 45 |  |  |  | 0.05165
| 848.31 | 11 | 2.67 | 29.37 | 29.3581 | 51 | -21.6419 | M4 | 20.8279 | 3.29 | 24.1179 | 45 | -20.8821 | M4 | 0.75977 | -1.4422
| 848.52 | 16.7 | 2.81 | 46.927 | 33.4285 | 51 | -17.5715 | M4 | 24.4543 | 3.28 | 27.7343 | 45 | -17.2657 | M4 | 0.30587 | 2.1843
| 848.31 | 11.47 | 2.79 | 32.0013 | 30.1034 | 51 | -20.8966 | M4 | 21.1913 | 3.2 | 24.3913 | 45 | -20.6087 | M4 | 0.28792 | -1.0788
| 848.31 | 13.5 | 2.58 | 34.83 | 30.8391 | 51 | -20.1609 | M4 | 22.6067 |  |  | 45 |  |  |  | 0.33664
PCS CDMA Full Rate | 1851.25 | 20.1 | 0.99 | 19.899 | 25.9766 | 41 | -15.0234 | M4 | 26.0639 | -19.38 | 6.6839 | 35 | -28.3161 | M4 | -13.2927 | -0.32759
| 1851.25 | 21.2 | 0.99 | 20.988 | 26.4394 | 41 | -14.5606 | M4 | 26.5267 | -18.5 | 8.0267 | 35 | -26.9733 | M4 | -12.4127 | 0.13521
| 1851.25 | 21.34 | 1.03 | 21.9802 | 26.8406 | 41 | -14.1594 | M4 | 26.5839 | -18.8 | 7.7839 | 35 | -27.2161 | M4 | -13.0567 | 0.19238
| 1851.25 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -26.3915
| 1880 | 20.9 | 0.99 | 20.691 | 26.3156 | 41 | -14.6844 | M4 | 26.4029 | -19.38 | 7.0229 | 35 | -27.9771 | M4 | -13.2927 | -0.34731
| 1880 | 21.6 | 0.99 | 21.384 | 26.6018 | 41 | -14.3982 | M4 | 26.6891 | -18.5 | 8.1891 | 35 | -26.8109 | M4 | -12.4127 | -0.061158
| 1880 | 22.8 | 1.03 | 23.484 | 27.4154 | 41 | -13.5846 | M4 | 27.1587 | -19.8 | 7.3587 | 35 | -27.6413 | M4 | -14.0567 | 0.40846
| 1880 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -26.7502
| 1908.75 | 22.1 | 0.99 | 21.879 | 26.8005 | 41 | -14.1995 | M4 | 26.8878 | -19.38 | 7.5078 | 35 | -27.4922 | M4 | -13.2927 | -0.12251
| 1908.5 | 22 | 0.99 | 21.78 | 26.7612 | 41 | -14.2388 | M4 | 26.8485 | -18.5 | 8.3485 | 35 | -26.6515 | M4 | -12.4127 | -0.1619
| 1908.75 | 23.16 | 1.03 | 23.8548 | 27.5515 | 41 | -13.4485 | M4 | 27.2948 | -18.1 | 9.1948 | 35 | -25.8052 | M4 | -12.3567 | 0.28441
| 1908.75 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -27.0104
PCS CDMA 1/8 Rate | 1851.25 | 7.09 | 2.765 | 19.6039 | 25.8468 | 41 | -15.1532 | M4 | 17.0129 | 3.24 | 20.2529 | 35 | -14.7471 | M4 | 0.4061 | -0.81036
| 1851.25 | 8.14 | 2.81 | 22.8734 | 27.1866 | 41 | -13.8134 | M4 | 18.2125 | 3.18 | 21.3925 | 35 | -13.6075 | M4 | 0.20587 | 0.3892
| 1851.25 | 8.17 | 2.77 | 22.6309 | 27.094 | 41 | -13.906 | M4 | 18.2444 | 3.2 | 21.4444 | 35 | -13.5556 | M4 | 0.3504 | 0.42116
| 1851.25 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -17.8233
| 1880 | 7.7 | 2.765 | 21.2905 | 26.5637 | 41 | -14.4363 | M4 | 17.7298 | 3.24 | 20.9698 | 35 | -14.0302 | M4 | 0.4061 | -0.88272
| 1880 | 8.78 | 2.81 | 24.6718 | 27.844 | 41 | -13.156 | M4 | 18.8699 | 3.18 | 22.0499 | 35 | -12.9501 | M4 | 0.20587 | 0.25735
| 1880 | 9.16 | 2.77 | 25.3732 | 28.0875 | 41 | -12.9125 | M4 | 19.2379 | 3.2 | 22.4379 | 35 | -12.5621 | M4 | 0.3504 | 0.62537
| 1880 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -18.6125
| 1908.75 | 8.21 | 2.765 | 22.7007 | 27.1208 | 41 | -13.8792 | M4 | 18.2869 | 3.24 | 21.5269 | 35 | -13.4731 | M4 | 0.4061 | -0.32986
| 1908.5 | 8.95 | 2.81 | 25.1495 | 28.0106 | 41 | -12.9894 | M4 | 19.0365 | 3.18 | 22.2165 | 35 | -12.7835 | M4 | 0.20587 | 0.41974
| 1908.75 | 8.44 | 2.77 | 23.3788 | 27.3764 | 41 | -13.6236 | M4 | 18.5268 | 3.2 | 21.7268 | 35 | -13.2732 | M4 | 0.3504 | -0.089875
| 1908.75 |  |  |  |  | 41 | [5 empty] | 35 |  |  |  | -18.6167

## Phone 7_GSM & WCDMA
Phone Performance Comparison Data
|  | GSM |  | WCDMA
Phone model: | SGH-i637 | CMU200 setup |  | CMU200 setup
Manufacturer | Samsung | Traffic Mode | Full Rate Version 2 | DCH Type | Speech
|  | Bit Stream | Speech Codec/ Handset Low | Voice Source | Speechcodec Low
IMEI | 35952202.051937.702 |  |  | Codec | AMR 12.2 kbps
FCC ID | A3LSGHI637 | [6 empty] | ;= reduced by the exclusion (one site only reporting this)
|  | Measured Avg. E-Field (V/m) | ANSI C63.19-2007 Calculations | [5 empty] | ANSI C63.19 Draft 1.13 Calculations | [5 empty] | 2007 v.s. Draft 1.13 Delta
|  |  | Measured PMF | Peak E-Field |  | M3 Limit | Margin to M3 Limit | M-Rating (2007) | Measured Avg. E-Field | Measured MIF | RF Audio Interference Level | M3 Limit | Margin to M3 Limit | M-Rating (draft 1.13)
Units |  | V/m | linear | V/m | dBV/m | dBV/m | dB |  | dBV/m | dB | dBV/m | dBV/m | dB |  | dB | Avg. E-Field deviation from the mean (dB)
Mode & Band | Frequency (MHz) | a | b | c = a * b | d = 20 log c | e | f = d - e |  | h = 20 log a | g | i = g + h | j | k = i - j |  | l = k - f
GSM 850 | 824.2 | 60.4 | 2.829 | 170.8716 | 44.6534 | 48.5 | -3.8466 | M3 | 35.6207 | 3.5 | 39.1207 | 45 | -5.8793 | M4 | -2.0327 | -0.15235
| 824.2 | 62.8 | 2.829 | 177.6612 | 44.9919 | 48.5 | -3.5081 | M3 | 35.9592 | 3.41 | 39.3692 | 45 | -5.6308 | M4 | -2.1227 | 0.18611
| 824.2 | 60.22 | 2.86 | 172.2292 | 44.7221 | 48.5 | -3.7779 | M3 | 35.5948 | 3.5 | 39.0948 | 45 | -5.9052 | M4 | -2.1273 | -0.17827
| 824.2 | 62.5 | 3 | 187.5 | 45.46 | 48.5 | -3.04 | M3 | 35.9176 |  |  | 45 |  |  |  | 0.14451
| 836.8 | 55.6 | 2.829 | 157.2924 | 43.9342 | 48.5 | -4.5658 | M3 | 34.9015 | 3.5 | 38.4015 | 45 | -6.5985 | M4 | -2.0327 | -1.1331
| 836.8 | 69.3 | 2.829 | 196.0497 | 45.8473 | 48.5 | -2.6527 | M3 | 36.8147 | 3.41 | 40.2247 | 45 | -4.7753 | M3 | -2.1227 | 0.78009
| 836.8 | 62.1 | 2.86 | 177.606 | 44.9892 | 48.5 | -3.5108 | M3 | 35.8618 | 3.5 | 39.3618 | 45 | -5.6382 | M4 | -2.1273 | -0.17274
| 836.8 | 67.3 | 3.01 | 202.573 | 46.1316 | 48.5 | -2.3684 | M3 | 36.5603 |  |  | 45 |  |  |  | 0.52573
| 848.8 | 55.4 | 2.829 | 156.7266 | 43.9029 | 48.5 | -4.5971 | M3 | 34.8702 | 3.5 | 38.3702 | 45 | -6.6298 | M4 | -2.0327 | -0.71185
| 848.8 | 65.6 | 2.829 | 185.5824 | 45.3707 | 48.5 | -3.1293 | M3 | 36.3381 | 3.41 | 39.7481 | 45 | -5.2519 | M4 | -2.1227 | 0.75603
| 848.8 | 57.93 | 2.86 | 165.6798 | 44.3854 | 48.5 | -4.1146 | M3 | 35.2581 | 3.5 | 38.7581 | 45 | -6.2419 | M4 | -2.1273 | -0.32397
| 848.8 | 62.1 | 3 | 186.3 | 45.4043 | 48.5 | -3.0957 | M3 | 35.8618 |  |  | 45 |  |  |  | 0.27979
GSM 1900 | 1850.2 | 18.9 | 2.79 | 52.731 | 34.4413 | 38.5 | -4.0587 | M3 | 25.5292 | 3.29 | 28.8192 | 35 | -6.1808 | M4 | -2.1221 | -0.9286
| 1850.2 | 18.5 | 2.79 | 51.615 | 34.2555 | 38.5 | -4.2445 | M3 | 25.3434 | 3.39 | 28.7334 | 35 | -6.2666 | M4 | -2.0221 | -1.1144
| 1850.2 | 27.57 | 2.84 | 78.2988 | 37.8751 | 38.5 | -0.6249 | M3 | 28.8087 | 3.5 | 32.3087 | 35 | -2.6913 | M3 | -2.0664 | 2.3509
| 1850.2 | 20.3 | 2.87 | 58.261 | 35.3076 | 38.5 | -3.1924 | M3 | 26.1499 |  |  | 35 |  |  |  | -0.30791
| 1880 | 16.7 | 2.79 | 46.593 | 33.3664 | 38.5 | -5.1336 | M4 | 24.4543 | 3.29 | 27.7443 | 35 | -7.2557 | M4 | -2.1221 | -1.3178
| 1880 | 17.2 | 2.79 | 47.988 | 33.6227 | 38.5 | -4.8773 | M3 | 24.7106 | 3.39 | 28.1006 | 35 | -6.8994 | M4 | -2.0221 | -1.0615
| 1880 | 26.01 | 2.84 | 73.8684 | 37.3692 | 38.5 | -1.1308 | M3 | 28.3028 | 3.5 | 31.8028 | 35 | -3.1972 | M3 | -2.0664 | 2.5307
| 1880 | 19.1 | 2.86 | 54.626 | 34.748 | 38.5 | -3.752 | M3 | 25.6207 |  |  | 35 |  |  |  | -0.15143
| 1909.8 | 13.9 | 2.79 | 38.781 | 31.7724 | 38.5 | -6.7276 | M4 | 22.8603 | 3.29 | 26.1503 | 35 | -8.8497 | M4 | -2.1221 | -1.2157
| 1909.8 | 14.7 | 2.79 | 41.013 | 32.2584 | 38.5 | -6.2416 | M4 | 23.3463 | 3.39 | 26.7363 | 35 | -8.2637 | M4 | -2.0221 | -0.72965
| 1909.8 | 20.24 | 2.84 | 57.4816 | 35.1906 | 38.5 | -3.3094 | M3 | 26.1242 | 3.5 | 29.6242 | 35 | -5.3758 | M4 | -2.0664 | 2.0482
| 1909.8 | 15.8 | 2.84 | 44.872 | 33.0395 | 38.5 | -5.4605 | M4 | 23.9731 |  |  | 35 |  |  |  | -0.10286
WCDMA 850 | 826.4 | 52.8 | 0.961 | 50.7408 | 34.1071 | 51 | -16.8929 | M4 | 34.4527 | -14.4 | 20.0527 | 45 | -24.9473 | M4 | -8.0545 | -0.061418
| 826.4 | 55.7 | 0.967 | 53.8619 | 34.6256 | 51 | -16.3744 | M4 | 34.9171 | -20.71 | 14.2071 | 45 | -30.7929 | M4 | -14.4185 | 0.40301
| 826.4 | 50.38 | 0.99 | 49.8762 | 33.9579 | 51 | -17.0421 | M4 | 34.0452 | -21.4 | 12.6452 | 45 | -32.3548 | M4 | -15.3127 | -0.46893
| 826.4 | 53.96 | 0.98 | 52.8808 | 34.466 | 51 | -16.534 | M4 | 34.6414 |  |  | 45 |  |  |  | 0.12734
| 836.4 | 45.1 | 0.961 | 43.3411 | 32.738 | 51 | -18.262 | M4 | 33.0835 | -14.4 | 18.6835 | 45 | -26.3165 | M4 | -8.0545 | -1.1496
| 836.4 | 59.5 | 0.967 | 57.5365 | 35.1989 | 51 | -15.8011 | M4 | 35.4903 | -20.71 | 14.7803 | 45 | -30.2197 | M4 | -14.4185 | 1.2572
| 836.4 | 48.84 | 0.99 | 48.3516 | 33.6882 | 51 | -17.3118 | M4 | 33.7755 | -25.2 | 8.5755 | 45 | -36.4245 | M4 | -19.1127 | -0.45766
| 836.4 | 53.6 | 0.98 | 52.528 | 34.4078 | 51 | -16.5922 | M4 | 34.5833 |  |  | 45 |  |  |  | 0.35013
| 846.6 | 53.9 | 0.961 | 51.7979 | 34.2862 | 51 | -16.7138 | M4 | 34.6318 | -14.4 | 20.2318 | 45 | -24.7682 | M4 | -8.0545 | -0.75936
| 846.6 | 68.4 | 0.967 | 66.1428 | 36.4097 | 51 | -14.5903 | M4 | 36.7011 | -20.71 | 15.9911 | 45 | -29.0089 | M4 | -14.4185 | 1.31
| 846.6 | 55.14 | 0.99 | 54.5886 | 34.742 | 51 | -16.258 | M4 | 34.8293 | -25.1 | 9.7293 | 45 | -35.2707 | M4 | -19.0127 | -0.5618
| 846.6 | 58.9 | 0.98 | 57.722 | 35.2268 | 51 | -15.7732 | M4 | 35.4023 |  |  | 45 |  |  |  | 0.011171
WCDMA 1900 | 1852.4 | 26.4 | 0.972 | 25.6608 | 28.1854 | 41 | -12.8146 | M4 | 28.4321 | -14.4 | 14.0321 | 35 | -20.9679 | M4 | -8.1533 | -0.65529
| 1852.4 | 25.8 | 0.989 | 25.5162 | 28.1363 | 41 | -12.8637 | M4 | 28.2324 | -20.1 | 8.1324 | 35 | -26.8676 | M4 | -14.0039 | -0.85497
| 1852.4 | 34.69 | 1.01 | 35.0369 | 30.8905 | 41 | -10.1095 | M4 | 30.8041 | -25.4 | 5.4041 | 35 | -29.5959 | M4 | -19.4864 | 1.7167
| 1852.4 | 27.8 | 0.99 | 27.522 | 28.7936 | 41 | -12.2064 | M4 | 28.8809 |  |  | 35 |  |  |  | -0.20647
| 1880 | 24.1 | 0.972 | 23.4252 | 27.3937 | 41 | -13.6063 | M4 | 27.6403 | -14.4 | 13.2403 | 35 | -21.7597 | M4 | -8.1533 | -0.64833
| 1880 | 23.8 | 0.989 | 23.5382 | 27.4355 | 41 | -13.5645 | M4 | 27.5315 | -20.1 | 7.4315 | 35 | -27.5685 | M4 | -14.0039 | -0.75713
| 1880 | 32.49 | 1.01 | 32.8149 | 30.3214 | 41 | -10.6786 | M4 | 30.235 | -25.3 | 4.935 | 35 | -30.065 | M4 | -19.3864 | 1.9463
| 1880 | 24.4 | 1 | 24.4 | 27.7478 | 41 | -13.2522 | M4 | 27.7478 |  |  | 35 |  |  |  | -0.54087
| 1907.6 | 20 | 0.972 | 19.44 | 25.7739 | 41 | -15.2261 | M4 | 26.0206 | -14.4 | 11.6206 | 35 | -23.3794 | M4 | -8.1533 | -1.1137
| 1907.6 | 22 | 0.989 | 21.758 | 26.7524 | 41 | -14.2476 | M4 | 26.8485 | -20.1 | 6.7485 | 35 | -28.2515 | M4 | -14.0039 | -0.28585
| 1907.6 | 27.48 | 1.01 | 27.7548 | 28.8668 | 41 | -12.1332 | M4 | 28.7803 | -23.4 | 5.3803 | 35 | -29.6197 | M4 | -17.4864 | 1.646
| 1907.6 | 22.1 | 1 | 22.1 | 26.8878 | 41 | -14.1122 | M4 | 26.8878 |  |  | 35 |  |  |  | -0.24646

## MIF
[9 empty] | GSM | CDMA 1/8 Rate | iDEN | CDMA Full Rate | WCDMA
| phone 1 | phone 2 | phone 3 | phone 4 | phone 5 | phone 6 | phone 7
GSM 850 | phone 1: 3.5; phone 2: 3.5; phone 4: 3.47; phone 7: 3.5 |  | 3.27 | 3.15 | 1.44 | -19.8 | -25.4
| phone 1: 3.52; phone 2: 3.55; phone 4: 3.49; phone 7: 3.41 |  | 3.29 | 3.18 | 1.51 | -19.7 | -25.3
| phone 1: 3.5 |  | 3.3 | 3.2 |  | -19.6 | -25.2
[9 empty] | 3.3 | 3.21 |  | -19.6 | -25.1
GSM 1900 | phone 1: 3.3; phone 2: 3.3; phone 4: 3.27; phone 7: 3.29 |  | 3.38 | 3.23 |  | -19.4 | -23.4
| phone 1: 3.52; phone 2: 3.54; phone 4: 3.38; phone 7: 3.39 |  | 3.39 | 3.24 |  | -19.4 | -21.8
[9 empty] | 3.41 | 3.24 |  | -19.4 | -21.5
[9 empty] | 3.47 | 3.24 |  | -19.38 | -21.4
WCDMA 850 | phone 1: -17; phone 4: -13.8; phone 7: -14.4 |  | 3.49 | 3.26 |  | -19.36 | -21.1
| phone 1: -19.94; phone 4: -20.01; phone 7: -20.71 |  | 3.5 | 3.27 |  | -19.3 | -21
| phone 1: -20.7; phone 4: -20.3; phone 7: -21.4 |  | 3.5 | 3.28 |  | -19.3 | -20.8
[9 empty] | 3.5 | 3.29 |  | -19.3 | -20.71
| phone 1: -21.1; phone 4: -20.2; phone 7: -25.2 |  | 3.5 |  |  | -19.3 | -20.7
| phone 1: -21; phone 4: -21.5; phone 7: -25.1 |  | 3.52 |  |  | -19.25 | -20.3
WCDMA 1900 | phone 1: -15.1; phone 4: -13.9; phone 7: -14.4 |  | 3.52 |  |  | -19.2 | -20.3
| phone 1: -20.15; phone 4: -20.13; phone 7: -20.1 |  | 3.54 |  |  | -19.2 | -20.2
| phone 1: -19.9; phone 4: -20.3; phone 7: -25.4 |  | 3.55 |  |  | -19.2 | -20.15
[12 empty] | -19.16 | -20.13
| phone 1: -19.9; phone 4: -20.8; phone 7: -25.3 |  |  |  |  | -19.13 | -20.1
| phone 1: -20.1; phone 4: -21.8; phone 7: -23.4 |  |  |  |  | -19.12 | -20.1
Cellular CDMA Full Rate | phone 2: -19.25; phone 5: -18.3; phone 6: -19.12 |  |  |  |  | -19.12 | -20.01
| phone 2: -19.13; phone 6: -19.12 |  |  |  |  | -19.1 | -19.94
| phone 2: -19.3; phone 5: -19.4; phone 6: -19.1 |  |  |  |  | -19.1 | -19.9
[12 empty] | -19.1 | -19.9
| phone 2: -19.3; phone 5: -19.1; phone 6: -19.1 |  |  |  |  | -18.8 | -17
| phone 2: -19.3; phone 5: -19.2; phone 6: -19.2 |  |  |  |  | -18.5 | -15.1
Cellular CDMA 1/8 Rate | phone 2: 3.23; phone 5: 3.27; phone 6: 3.29 |  |  |  |  | -18.5 | -14.4
| phone 2: 3.21; phone 6: 3.28 |  |  |  |  | -18.3 | -14.4
| phone 2: 3.2 |  |  |  |  | -18.1 | -13.9
-13.8
PCS CDMA Full Rate | phone 2: -19.16; phone 5: -19.36; phone 6: -19.38
|  | -19.3 |  |  | -18.5 | -18.5
|  | -19.2 |  |  | -19.4 | -18.8
|  | -19.7 |  |  | -19.6 | -19.8
|  | -19.4 |  |  | -19.6 | -18.1
PCS CDMA 1/8 Rate | phone 2: 3.26; phone 5: 3.24; phone 6: 3.24
| phone 2: 3.24; phone 5: 3.15; phone 6: 3.18 |  | GSM | CDMA 1/8 Rate | iDEN | CDMA Full Rate | WCDMA
[8 empty] | mean | 3.4371 | 3.2325 | 1.475 | -19.1628 | -20.1013
iDEN 800 | phone 3: 1.51 | std. dev. | 0.096874 | 0.041588 | 0.049497 | 0.39277 | 3.2405
iDEN 900 | phone 3: 1.44 | max | 3.55 | 3.29 | 1.51 | -18.1 | -13.8
[8 empty] | min | 3.27 | 3.15 | 1.44 | -19.8 | -25.4
[8 empty] | max-min | 0.28 | 0.14 | 0.07 | 1.7 | 11.6
| phone 7: less WCDMA MIF outliers | mean |  |  |  |  | -20.5284
[8 empty] | std. dev. |  |  |  |  | 0.59123
[8 empty] | max |  |  |  |  | -19.9
[8 empty] | min |  |  |  |  | -21.8
[8 empty] | max-min |  |  |  |  | 1.9

## delta dB
[9 empty] | GSM | CDMA 1/8 Rate | iDEN | CDMA Full Rate | WCDMA
| phone 1 | phone 2 | phone 3 | phone 4 | phone 5 | phone 6 | phone 7
GSM 850 | phone 1: -2.0327; phone 2: -2.0327; phone 4: -2.0627; phone 7: -2.0327 |  | -2.2874 | 0.0044763 | -2.9389 | -14.0567 | -19.4864
| phone 1: -2.1376; phone 2: -2.2874; phone 4: -2.0457; phone 7: -2.1227 |  | -2.2677 | 0.17363 | -2.1601 | -13.9567 | -19.3864
| phone 1: -2.1273 |  | -2.1421 | 0.20587 |  | -13.8567 | -19.1127
[9 empty] | -2.1376 | 0.28792 |  | -13.8567 | -19.0127
[9 empty] | -2.1273 | 0.30587 |  | -13.6567 | -17.4864
[9 empty] | -2.125 | 0.3504 |  | -13.6567 | -15.8864
[9 empty] | -2.1227 | 0.4061 |  | -13.4567 | -15.4127
[9 empty] | -2.1221 | 0.4061 |  | -13.3 | -15.3127
[9 empty] | -2.1121 | 0.4104 |  | -13.2927 | -15.0127
[9 empty] | -2.1121 | 0.6473 |  | -13.2727 | -14.9127
[9 empty] | -2.0664 | 0.69977 |  | -13.1354 | -14.8864
[9 empty] | -2.0627 | 0.73977 |  | -13.0727 | -14.6127
GSM 1900 | phone 1: -2.1121; phone 2: -2.1121; phone 4: -2.1421; phone 7: -2.1221 |  | -2.0457 | 0.75977 |  | -13.0567 | -14.4185
| phone 1: -1.8921; phone 2: -2.2677; phone 4: -2.125; phone 7: -2.0221 |  | -2.0327 |  |  | -13.0354 | -14.3867
| phone 1: -2.0664 |  | -2.0327 |  |  | -13.0354 | -14.3864
[9 empty] | -2.0327 |  |  | -13.0354 | -14.2127
[9 empty] | -2.0221 |  |  | -12.9354 | -14.1864
[9 empty] | -1.8921 |  |  | -12.9354 | -14.182
[12 empty] | -12.8354 | -14.1127
[12 empty] | -12.8354 | -14.0039
[12 empty] | -12.8354 | -13.9864
[12 empty] | -12.7754 | -13.9864
[12 empty] | -12.6197 | -13.4026
[12 empty] | -12.5864 | -13.3308
WCDMA 850 | phone 1: -10.6545; phone 4: -7.4545; phone 7: -8.0545 |  |  |  |  | -12.471 | -10.6545
| phone 1: -13.4026; phone 4: -14.182; phone 7: -14.4185 |  |  |  |  | -12.471 | -8.8533
| phone 1: -14.6127; phone 4: -14.2127; phone 7: -15.3127 |  |  |  |  | -12.4127 | -8.1533
[12 empty] | -12.3567 | -8.0545
[12 empty] | -11.651 | -7.6533
-7.4545
| -15.0127 |  |  | -14.1127 |  |  | -19.1127
| -14.9127 |  |  | -15.4127 |  |  | -19.0127
WCDMA 1900 | phone 1: -8.8533; phone 4: -7.6533; phone 7: -8.1533
| -13.3308 |  |  | -14.3867 |  |  | -14.0039
| -13.9864 |  |  | -14.3864 |  |  | -19.4864
| -13.9864 |  |  | -14.8864 |  |  | -19.3864
| -14.1864 |  |  | -15.8864 |  |  | -17.4864
Cellular CDMA Full Rate | phone 2: -12.6197; phone 5: -11.651; phone 6: -12.471
|  | -12.7754 |  |  |  | -12.471
|  | -13.0354 |  |  | -13.1354 | -12.8354
|  | -13.0354 |  |  | -12.8354 | -12.8354
|  | -13.0354 |  |  | -12.9354 | -12.9354
Cellular CDMA 1/8 Rate | phone 2: 0.69977; phone 5: 0.73977; phone 6: 0.75977
|  | 0.6473 |  |  |  | 0.30587
0.28792
PCS CDMA Full Rate | phone 2: -13.0727; phone 5: -13.2727; phone 6: -13.2927
|  | -13.3 |  |  |  | -12.4127
|  | -13.4567 |  |  | -13.6567 | -13.0567
-12.5864
|  | -13.9567 |  |  | -13.8567 | -14.0567
|  | -13.6567 |  |  | -13.8567 | -12.3567
PCS CDMA 1/8 Rate | phone 2: 0.4104; phone 5: 0.4061; phone 6: 0.4061 |  | GSM | CDMA 1/8 Rate | iDEN | CDMA Full Rate | WCDMA
|  | 0.17363 |  |  |  | 0.20587
| phone 2: 0.3504 | mean | -2.0968 | 0.41518 | -2.5495 | -13.0502 | -13.998
[8 empty] | std. dev. | 0.089854 | 0.23494 | 0.55066 | 0.54684 | 3.3455
[8 empty] | max | -1.8921 | 0.75977 | -2.1601 | -11.651 | -7.4545
| phone 5: 0.0044763 | min | -2.2874 | 0.0044763 | -2.9389 | -14.0567 | -19.4864
[8 empty] | max-min | 0.39527 | 0.7553 | 0.77875 | 2.4058 | 12.032
| phone 7: less WDMA MIF outliers | mean |  |  |  |  | -14.4543
[8 empty] | std. dev. |  |  |  |  | 0.65746
[8 empty] | max |  |  |  |  | -13.3308
[8 empty] | min |  |  |  |  | -15.8864
iDEN 800 | phone 3: -2.9389 | max-min |  |  |  |  | 2.5556
iDEN 900 | phone 3: -2.1601

## RF consistency
[9 empty] | low band |  |  |  | high band |  |  |  |  | low and high band combined
[9 empty] | GSM | CDMA 1/8 Rate | CDMA Full Rate | WCDMA | GSM | CDMA 1/8 Rate | CDMA Full Rate | WCDMA |  | GSM | CDMA 1/8 Rate | CDMA Full Rate | WCDMA
| phone 1 | phone 2 | phone 3 | phone 4 closed | phone 4 open | phone 5 | phone 6 | phone 7
GSM 850 | -0.012207 | -0.57817 |  | -0.45346 | -0.39769 |  |  | -0.15235 | -1.3348 | -1.4422 | -1.0867 | -1.3416 | -1.3178 | -1.2192 | -0.97152 | -1.2934 |  | -1.3348 | -1.4422 | -1.0867 | -1.3416
| 0.25102 | 1.1488 |  | 0.3503 | 1.096 |  |  | 0.18611 | -1.334 | -1.1804 | -0.8547 | -1.2105 | -1.2349 | -1.2188 | -0.85072 | -1.2567 |  | -1.334 | -1.2192 | -0.97152 | -1.2934
| -0.30053 | -0.76535 |  | -0.069868 | -0.53593 |  |  | -0.17827 | -1.2746 | -1.0864 | -0.80397 | -1.1733 | -1.2157 | -1.0101 | -0.48351 | -1.1137 |  | -1.3178 | -1.2188 | -0.8547 | -1.2567
| 0.061716 | 0.19473 |  | 0.17303 | -0.1624 |  |  | 0.14451 | -1.1699 | -1.0788 | -0.67007 | -1.1496 | -1.2013 | -0.88272 | -0.43629 | -1.0376 |  | -1.2746 | -1.1804 | -0.85072 | -1.2105
| -0.82195 | -1.3348 |  | -1.2746 | -1.1699 |  |  | -1.1331 | -1.1331 | -0.84 | -0.66645 | -1.1133 | -1.1144 | -0.81036 | -0.43453 | -1.0029 |  | -1.2349 | -1.0864 | -0.80397 | -1.1733
| 0.50202 | 0.80064 |  | 0.84461 | 1.2652 |  |  | 0.78009 | -1.006 | -0.80216 | -0.61239 | -1.029 | -1.0615 | -0.80751 | -0.34731 | -0.91774 |  | -1.2157 | -1.0788 | -0.67007 | -1.1496
| -0.24178 | 0.037143 |  | -0.30654 | -0.14163 |  |  | -0.17274 | -0.9347 | -0.63077 | -0.59648 | -0.90436 | -1.0268 | -0.79799 | -0.32759 | -0.85497 |  | -1.2013 | -1.0101 | -0.66645 | -1.1137
| 0.56172 | 0.49707 |  | 0.73648 | 0.046401 |  |  | 0.52573 | -0.82195 | -0.61317 | -0.5667 | -0.8023 | -1.0253 | -0.72892 | -0.24256 | -0.78391 |  | -1.1699 | -0.88272 | -0.61239 | -1.1133
| -0.74716 | -1.006 |  | -1.334 | -0.9347 |  |  | -0.71185 | -0.76535 | -0.40616 | -0.41002 | -0.75936 | -0.9286 | -0.62566 | -0.19117 | -0.75713 |  | -1.1331 | -0.84 | -0.59648 | -1.0376
| 0.55409 | 0.67206 |  | 1.0426 | 1.16 |  |  | 0.75603 | -0.74716 | -0.37687 | -0.2489 | -0.5618 | -0.88611 | -0.32986 | -0.1619 | -0.74734 |  | -1.1144 | -0.81036 | -0.5667 | -1.029
| 0.13634 | 0.13212 |  | -0.053231 | 0.13648 |  |  | -0.32397 | -0.71185 | -0.16809 | -0.22255 | -0.53501 | -0.87637 | -0.30203 | -0.15649 | -0.71572 |  | -1.0615 | -0.80751 | -0.48351 | -1.0029
| 0.056729 | 0.20179 |  | 0.34461 | -0.36179 |  |  | 0.27979 | -0.57817 | -0.15645 | -0.2105 | -0.46893 | -0.85639 | -0.28407 | -0.12251 | -0.69574 |  | -1.0268 | -0.80216 | -0.43629 | -0.91774
GSM 1900 | -0.0142 | -0.11037 |  | -0.53243 | -0.78598 |  |  | -0.9286 | -0.53593 | -0.086397 | -0.20887 | -0.45766 | -0.84425 | -0.127 | -0.061158 | -0.65529 |  | -1.0253 | -0.79799 | -0.43453 | -0.90436
| -0.64171 | -0.66582 |  | -0.75345 | -1.0253 |  |  | -1.1144 | -0.45346 | -0.039303 | -0.075531 | -0.43603 | -0.79124 | -0.089875 | -0.0073355 | -0.64833 |  | -1.006 | -0.72892 | -0.41002 | -0.85497
| 1.2976 | 0.69649 |  | 1.9947 | 2.0485 |  |  | 2.3509 | -0.39769 | -0.035453 | -0.04492 | -0.40212 | -0.78598 | 0.13011 | -0.0040209 | -0.5536 |  | -0.9347 | -0.63077 | -0.34731 | -0.8023
| -0.64171 | 0.0797 |  | -0.7088 | -0.23721 |  |  | -0.30791 | -0.36179 | 0.035453 | -0.010181 | -0.38494 | -0.75345 | 0.17413 | 0.10046 | -0.54087 |  | -0.9286 | -0.62566 | -0.32759 | -0.78391
| -0.080926 | -0.34972 |  | -1.2013 | -0.84425 |  |  | -1.3178 | -0.32397 | 0.039303 | 0.04492 | -0.37232 | -0.72965 | 0.25735 | 0.13521 | -0.49735 |  | -0.88611 | -0.61317 | -0.2489 | -0.75936
| -0.85639 | -0.30941 |  | -1.0268 | -0.88611 |  |  | -1.0615 | -0.30654 | 0.05165 | 0.13115 | -0.31785 | -0.7088 | 0.3892 | 0.15636 | -0.46073 |  | -0.87637 | -0.40616 | -0.24256 | -0.75713
| 1.3651 | 0.65266 |  | 2.5496 | 1.9309 |  |  | 2.5307 | -0.30053 | 0.12242 | 0.1528 | -0.3113 | -0.68291 | 0.41974 | 0.19238 | -0.42679 |  | -0.85639 | -0.37687 | -0.22255 | -0.74734
| -0.42778 | 0.0064726 |  | -0.32154 | -0.20055 |  |  | -0.15143 | -0.24178 | 0.15852 | 0.20247 | -0.1954 | -0.66582 | 0.42116 | 0.25391 | -0.3989 |  | -0.84425 | -0.32986 | -0.2105 | -0.71572
| 0.42745 | -0.051567 |  | 0.10979 | -0.27189 |  |  | -1.2157 | -0.17827 | 0.16809 | 0.20887 | -0.14043 | -0.64171 | 0.47382 | 0.28441 | -0.37144 |  | -0.82195 | -0.30203 | -0.20887 | -0.69574
| -0.6114 | -0.5918 |  | -1.2349 | -0.87637 |  |  | -0.72965 | -0.17274 | 0.23574 | 0.2105 | -0.10445 | -0.64171 | 0.5861 | 0.32716 | -0.36578 |  | -0.79124 | -0.28407 | -0.19117 | -0.65529
| 0.97519 | 0.22761 |  | 1.7366 | 1.8312 |  |  | 2.0482 | -0.1624 | 0.33664 | 0.23335 | -0.079601 | -0.61144 | 0.62256 | 0.34766 | -0.29158 |  | -0.78598 | -0.16809 | -0.1619 | -0.64833
| -0.79124 | 0.41576 |  | -0.61144 | -0.68291 |  |  | -0.10286 | -0.15235 | 0.45932 | 0.23619 | -0.061418 | -0.6114 | 0.62537 | 0.40846 | -0.28585 |  | -0.76535 | -0.15645 | -0.15649 | -0.5618
WCDMA 850 | 0.16871 |  |  | -0.1954 | 0.35261 |  |  | -0.061418 | -0.14163 | 0.56258 | 0.63153 | -0.021874 | -0.5918 | 1.0531 | 0.43453 | -0.28185 |  | -0.75345 | -0.127 | -0.12251 | -0.5536
| 0.25528 |  |  | 0.37945 | 0.50025 |  |  | 0.40301 | -0.069868 | 0.649 | 0.74835 | 0.011171 | -0.53243 | 1.2188 | 0.85072 | -0.26258 |  | -0.74716 | -0.089875 | -0.075531 | -0.54087
| -0.40212 |  |  | -0.079601 | -0.53501 |  |  | -0.46893 | -0.053231 | 0.98113 | 0.76617 | 0.02084 | -0.42778 | 1.2192 | 1.3073 | -0.24646 |  | -0.72965 | -0.086397 | -0.061158 | -0.53501
| -0.021874 |  |  | -0.10445 | -0.31785 |  |  | 0.12734 | -0.012207 | 1.3105 | 0.97662 | 0.042466 | -0.34972 | 1.6435 |  | -0.20647 |  | -0.71185 | -0.039303 | -0.04492 | -0.49735
| -1.2105 |  |  | -0.90436 | -1.3416 |  |  | -1.1496 | 0.037143 | 1.648 | 1.299 | 0.068889 | -0.32154 |  |  | -0.096229 |  | -0.7088 | -0.035453 | -0.010181 | -0.46893
| 1.1681 |  |  | 1.3351 | 1.584 |  |  | 1.2572 | 0.046401 | 2.1843 | 1.447 | 0.12734 | -0.30941 |  |  | -0.080843 |  | -0.68291 | 0.035453 | -0.0073355 | -0.46073
| 0.042466 |  |  | -1.029 | -0.3113 |  |  | -0.45766 | 0.056729 |  |  | 0.16871 | -0.30791 |  |  | -0.0025894 |  | -0.66582 | 0.039303 | -0.0040209 | -0.45766
| -34.7077 |  |  | 0.59821 | 0.068889 |  |  | 0.35013 | 0.061716 |  |  | 0.25528 | -0.27189 |  |  | 0.097219 |  | -0.64171 | 0.05165 | 0.04492 | -0.43603
| -1.1733 |  |  | -1.1133 | -0.8023 |  |  | -0.75936 | 0.13212 |  |  | 0.35013 | -0.23721 |  |  | 0.23051 |  | -0.64171 | 0.12242 | 0.10046 | -0.42679
| 1.3137 |  |  | 1.4774 | 1.6107 |  |  | 1.31 | 0.13634 |  |  | 0.35261 | -0.20055 |  |  | 0.92033 |  | -0.61144 | 0.13011 | 0.13115 | -0.40212
| -0.14043 |  |  | -0.38494 | -0.37232 |  |  | -0.5618 | 0.13648 |  |  | 0.37945 | -0.15143 |  |  | 0.95586 |  | -0.6114 | 0.15852 | 0.13521 | -0.3989
| -36.012 |  |  | 0.02084 | -0.43603 |  |  | 0.011171 | 0.14451 |  |  | 0.40301 | -0.11037 |  |  | 1.0515 |  | -0.5918 | 0.16809 | 0.1528 | -0.38494
WCDMA 1900 | -0.0025894 |  |  | -0.36578 | -0.29158 |  |  | -0.65529 | 0.17303 |  |  | 0.50025 | -0.10286 |  |  | 1.0629 |  | -0.57817 | 0.17413 | 0.15636 | -0.37232
| -0.91774 |  |  | -0.28185 | -0.26258 |  |  | -0.85497 | 0.18611 |  |  | 0.59821 | -0.080926 |  |  | 1.1553 |  | -0.53593 | 0.23574 | 0.19238 | -0.37144
| 0.92033 |  |  | 1.6853 | 1.0515 |  |  | 1.7167 | 0.19473 |  |  | 1.1681 | -0.051567 |  |  | 1.646 |  | -0.53243 | 0.25735 | 0.20247 | -0.36578
| -31.5983 |  |  | -1.0376 | -0.49735 |  |  | -0.20647 | 0.20179 |  |  | 1.2572 | -0.0142 |  |  | 1.6853 |  | -0.45346 | 0.33664 | 0.20887 | -0.31785
| -0.37144 |  |  | -0.42679 | -0.096229 |  |  | -0.64833 | 0.25102 |  |  | 1.31 | 0.0064726 |  |  | 1.7167 |  | -0.42778 | 0.3892 | 0.2105 | -0.3113
| -0.78391 |  |  | -0.5536 | -0.3989 |  |  | -0.75713 | 0.27979 |  |  | 1.3137 | 0.0797 |  |  | 1.7277 |  | -0.39769 | 0.41974 | 0.23335 | -0.29158
| 1.1553 |  |  | 1.7277 | 0.95586 |  |  | 1.9463 | 0.34461 |  |  | 1.3351 | 0.10979 |  |  | 1.7795 |  | -0.36179 | 0.42116 | 0.23619 | -0.28585
| -31.1278 |  |  | -0.74734 | -0.46073 |  |  | -0.54087 | 0.3503 |  |  | 1.4774 | 0.22761 |  |  | 1.8752 |  | -0.34972 | 0.45932 | 0.25391 | -0.28185
| 0.23051 |  |  | 0.097219 | -0.080843 |  |  | -1.1137 | 0.49707 |  |  | 1.584 | 0.41576 |  |  | 1.9463 |  | -0.32397 | 0.47382 | 0.28441 | -0.26258
| -1.2934 |  |  | -0.71572 | -0.69574 |  |  | -0.28585 | 0.50202 |  |  | 1.6107 | 0.42745 |  |  |  |  | -0.32154 | 0.56258 | 0.32716 | -0.24646
| 1.0629 |  |  | 1.8752 | 1.7795 |  |  | 1.646 | 0.52573 |  |  |  | 0.65266 |  |  |  |  | -0.30941 | 0.5861 | 0.34766 | -0.20647
| -30.6011 |  |  | -1.2567 | -1.0029 |  |  | -0.24646 | 0.55409 |  |  |  | 0.69649 |  |  |  |  | -0.30791 | 0.62256 | 0.40846 | -0.1954
Cellular CDMA Full Rate |  | -0.5667 |  |  |  | 0.04492 | -0.66645 |  | 0.56172 |  |  |  | 0.97519 |  |  |  |  | -0.30654 | 0.62537 | 0.43453 | -0.14043
|  | 0.97662 |  |  |  | -36.7321 | -0.22255 |  | 0.67206 |  |  |  | 1.2976 |  |  |  |  | -0.30053 | 0.649 | 0.63153 | -0.10445
|  | -0.61239 |  |  |  | -0.04492 | 1.299 |  | 0.73648 |  |  |  | 1.3651 |  |  |  |  | -0.27189 | 0.98113 | 0.74835 | -0.096229
|  | 0.20247 |  |  |  | -36.7321 | -0.41002 |  | 0.75603 |  |  |  | 1.7366 |  |  |  |  | -0.24178 | 1.0531 | 0.76617 | -0.080843
|  | -0.67007 |  |  |  | -0.2105 | -0.8547 |  | 0.78009 |  |  |  | 1.8312 |  |  |  |  | -0.23721 | 1.2188 | 0.85072 | -0.079601
|  | 0.76617 |  |  |  | -35.6862 | 0.63153 |  | 0.80064 |  |  |  | 1.9309 |  |  |  |  | -0.20055 | 1.2192 | 0.97662 | -0.061418
|  | -0.2489 |  |  |  | 0.2105 | 0.23335 |  | 0.84461 |  |  |  | 1.9947 |  |  |  |  | -0.17827 | 1.3105 | 1.299 | -0.021874
|  | 0.1528 |  |  |  | -35.6862 | -0.010181 |  | 1.0426 |  |  |  | 2.0482 |  |  |  |  | -0.17274 | 1.6435 | 1.3073 | -0.0025894
|  | -0.80397 |  |  |  | -0.20887 | -1.0867 |  | 1.096 |  |  |  | 2.0485 |  |  |  |  | -0.1624 | 1.648 | 1.447 | 0.011171
|  | 0.74835 |  |  |  | -34.9845 | 1.447 |  | 1.1488 |  |  |  | 2.3509 |  |  |  |  | -0.15235 | 2.1843 |  | 0.02084
|  | -0.075531 |  |  |  | 0.20887 | -0.59648 |  | 1.16 |  |  |  | 2.5307 |  |  |  |  | -0.15143 |  |  | 0.042466
|  | 0.13115 |  |  |  | -34.9845 | 0.23619 |  | 1.2652 |  |  |  | 2.5496 |  |  |  |  | -0.14163 |  |  | 0.068889
Cellular CDMA 1/8 Rate |  | 0.23574 |  |  |  | 0.035453 | -1.1804 | [10 empty] | -0.11037 |  |  | 0.097219
|  | 0.98113 |  |  |  | -27.7832 | 0.56258 | [10 empty] | -0.10286 |  |  | 0.12734
|  | -0.84 |  |  |  | -0.035453 | 0.45932 | [10 empty] | -0.080926 |  |  | 0.16871
|  | -0.37687 |  |  |  | -27.7832 | 0.15852 | [10 empty] | -0.069868 |  |  | 0.23051
|  | -0.15645 |  |  |  | -0.16809 | -1.0864 | [10 empty] | -0.053231 |  |  | 0.25528
|  | 0.649 |  |  |  | -27.1342 | 1.648 | [10 empty] | -0.051567 |  |  | 0.35013
|  | -0.40616 |  |  |  | 0.16809 | -0.61317 | [10 empty] | -0.0142 |  |  | 0.35261
|  | -0.086397 |  |  |  | -27.1342 | 0.05165 | [10 empty] | -0.012207 |  |  | 0.37945
|  | -0.80216 |  |  |  | -0.039303 | -1.4422 | [10 empty] | 0.0064726 |  |  | 0.40301
|  | 1.3105 |  |  |  | -26.4422 | 2.1843 | [10 empty] | 0.037143 |  |  | 0.50025
|  | -0.63077 |  |  |  | 0.039303 | -1.0788 | [10 empty] | 0.046401 |  |  | 0.59821
|  | 0.12242 |  |  |  | -26.4422 | 0.33664 | [10 empty] | 0.056729 |  |  | 0.92033
PCS CDMA Full Rate |  | -8.2401 |  |  |  | -0.43453 | -0.32759 | [10 empty] | 0.061716 |  |  | 0.95586
|  | -0.48351 |  |  |  | -30.6722 | 0.13521 | [10 empty] | 0.0797 |  |  | 1.0515
|  | 0.15636 |  |  |  | 0.43453 | 0.19238 | [10 empty] | 0.10979 |  |  | 1.0629
|  | 0.32716 |  |  |  | -30.6722 | -26.3915 | [10 empty] | 0.13212 |  |  | 1.1553
|  | -0.0073355 |  |  |  | -0.43629 | -0.34731 | [10 empty] | 0.13634 |  |  | 1.1681
|  | -0.24256 |  |  |  | -0.97152 | -0.061158 | [10 empty] | 0.13648 |  |  | 1.2572
|  | -0.0040209 |  |  |  | 1.3073 | 0.40846 | [10 empty] | 0.14451 |  |  | 1.31
|  | 0.25391 |  |  |  | 0.10046 | -26.7502 | [10 empty] | 0.17303 |  |  | 1.3137
|  | 0.34766 |  |  |  | -0.85072 | -0.12251 | [10 empty] | 0.18611 |  |  | 1.3351
|  | -28.15 |  |  |  | -28.3823 | -0.1619 | [10 empty] | 0.19473 |  |  | 1.4774
|  | -0.19117 |  |  |  | 0.85072 | 0.28441 | [10 empty] | 0.20179 |  |  | 1.584
|  | -0.15649 |  |  |  | -28.3823 | -27.0104 | [10 empty] | 0.22761 |  |  | 1.6107
PCS CDMA 1/8 Rate |  | -0.62566 |  |  |  | -1.2192 | -0.81036 | [10 empty] | 0.25102 |  |  | 1.646
|  | 0.62256 |  |  |  | -22.2036 | 0.3892 | [10 empty] | 0.27979 |  |  | 1.6853
|  | -0.127 |  |  |  | 1.2192 | 0.42116 | [10 empty] | 0.34461 |  |  | 1.7167
|  | 0.13011 |  |  |  | -22.2036 | -17.8233 | [10 empty] | 0.3503 |  |  | 1.7277
|  | 1.0531 |  |  |  | -1.0101 | -0.88272 | [10 empty] | 0.41576 |  |  | 1.7795
|  | 0.47382 |  |  |  | -0.80751 | 0.25735 | [10 empty] | 0.42745 |  |  | 1.8752
|  | -0.79799 |  |  |  | 1.6435 | 0.62537 | [10 empty] | 0.49707 |  |  | 1.9463
|  | -0.72892 |  |  |  | 0.17413 | -18.6125 | [10 empty] | 0.50202
|  | -0.30203 |  |  |  | -1.2188 | -0.32986 | [10 empty] | 0.52573
|  | -20.0016 |  |  |  | -19.4738 | 0.41974 | [10 empty] | 0.55409
|  | -0.28407 |  |  |  | 1.2188 | -0.089875 | [10 empty] | 0.56172
|  | 0.5861 |  |  |  | -19.4738 | -18.6167 | [10 empty] | 0.65266
0.67206
0.69649
0.73648
0.75603
0.78009
0.80064
0.84461
0.97519
1.0426
1.096
1.1488
1.16
1.2652
1.2976
1.3651
1.7366
1.8312
1.9309
1.9947
2.0482
2.0485
2.3509
2.5307
2.5496
[18 empty] | GSM | CDMA 1/8 Rate | CDMA Full Rate | WCDMA
[17 empty] | mean | 2.6645e-16 | -2.4501e-16 | -2.4931e-16 | -3.1233e-16
[17 empty] | std. dev. | 0.88776 | 0.79869 | 0.55687 | 0.88301
[17 empty] | max | 2.5496 | 2.1843 | 1.447 | 1.9463
[17 empty] | min | -1.3348 | -1.4422 | -1.0867 | -1.3416
[17 empty] | max-min | 3.8845 | 3.6265 | 2.5337 | 3.288

## graphs
Observation on WCDMA MIF measurements:
The highest 6 all came from the same lab on 3 different phones.
The lowest 5 all came from the same lab on 1 phone.
These slighly odd results carry on through to the second graph.
Phone setup differences?